## BSD&KL(Pdas)-Z8G4 CSD Menu
Service tools |  |  |  |  | BSD16 | [19 empty] | KL16-PDAS
System tools/utilities/diags | Tool License Token | Tools name | Tool overview | Upper Folder | A1 | A2 | A3 | C1 | C2 | C3 | M | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | Protection Level | Marks | Impact of reclassification to Indepedant Service Provider (Class A moved to Class C) | Difference | level | Comments | A1 | A2 | A3 | C1 | C2 | C3 | M | level | Comments
Error Logs (CSD menu)
System Error Log | ERROR_SYSTEM_LOG | System Error Log | [6 empty] | X | [6 empty] | X |  |  | 1900-01-01 | No protection to log files anyway. And it's using a viewlog command which may be widely used. |  |  | 2 | [5 empty] | X |  |  | 2
Software History Logs | ERROR_SOFTWARE_HISTORY_LOGS | Software History Log | [6 empty] | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 3 | [5 empty] | X |  |  | 3
Store Log | ERROR_STORE_LOG | Storelog |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 3 |  |  | X | [5 empty] | 3
System Browser(All feature) | ERROR_SYSTEM_BROWSER | System Browser (All features ) | [5 empty] | X | [6 empty] | X |  |  |  | 1899-12-31 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param. |  |  | 1 |  |  |  |  | X |  |  |  | 1
Tube Usage | ERROR_TUBE_USAGE | Tube Usage | [5 empty] | X | [6 empty] | X |  |  |  | 1899-12-31 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param. |  |  | 1 |  |  |  |  | X |  |  |  | 1
LogViewer | ERROR_LOG_VIEWER | LogViewer | [5 empty] | X | [6 empty] | X |  |  |  | 1900-01-02 |  |  |  | 3 | removed |  |  |  | X |  |  |  | 3
StartGFV | ERROR_START_GFV | StartGFV | [6 empty] | X | [6 empty] | X |  |  | 1899-12-31 | Technical difficulty: it's only a tar package installed to console. |  |  | 1 | [5 empty] | X |  |  | 1
Diagnostics (CSD menu)
LAN Diagnositics (System Communication) | DIAG_LAN | LAN Diagnostics |  | SystemCommunications |  |  |  |  | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 3 | [5 empty] | X |  |  | 3
CPU Processes (System Monitoring) | DIAG_CPU_PROCESSES | CPU Processes |  | SystemMonitoring |  |  |  |  | X | [6 empty] | X |  |  | 1899-12-31 | This is a Linux built-in command. No need to protect it. |  |  | 1 | [5 empty] | X |  |  | 1
kV & mA (X-ray Generation) | DIAG_KV_MA | kV & mA (X-ray) |  | X-RayGeneration |  | X | [6 empty] | X | [5 empty] | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs. |  |  | 2 |  |  | X | [5 empty] | 2
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI | Generator Tool - JEDI |  | X-RayGeneration |  | X |  |  | X |  | X |  | X |  |  | X |  | X | 1900-01-02 | Need to include KVMA in license |  |  | 4 |  |  | X |  |  | X |  | X | 4
Generator Tool – JEDI Filament Cal | CAL_GENERATOR_FILAMENT |  |  |  |  | X | [16 empty] | 4 |  |  | X | [5 empty] | 4
Generator Tool – JEDI HHS Scans | UTIL_HHS_SCANS | [7 empty] | X | [13 empty] | 4 | [5 empty] | X |  |  | 4
Generator Tool – JEDI HHS Report | UTIL_HHS_REPORT | [7 empty] | X | [13 empty] | 4 | [5 empty] | X |  |  | 4
Generator Tool – JEDI KV Diagnostics Inverter Short Circuit | DIAG_GENERATOR_TOOL_JEDI |  |  |  |  | X |  |  | X | [13 empty] | 4 |  |  | X |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics No Load KV | DIAG_GENERATOR_TOOL_JEDI |  |  |  |  | X |  |  | X | [13 empty] | 4 |  |  | X |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics No Load KV(Rotation) | [10 empty] | X | [11 empty] | 4 | [7 empty] | X | 4
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE | Collimator Aperture Test |  | X-RayGeneration |  |  |  | X | [6 empty] | X |  |  |  | 1900-01-02 |  |  |  | 4 |  |  |  |  | X |  |  |  | 4
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL | Collimator Functional |  | X-RayGeneration |  |  |  | X | [6 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs. |  |  | 2 |  |  |  |  | X |  |  |  | 2
Axial Functional (AxialControl) | DIAG_AXIAL_FUNCTIONAL | Axial Functional |  | AxialControl |  |  |  |  | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 3 | [5 empty] | X |  |  | 3
Gantry Service Balance (AxialControl) | DIAG_GANTRY_SERVICE_BALANCE |  |  | AxialControl |  | X | [15 empty] | Per SWRS.CSD.250, this was once defined as both Class A/C, don't know why later changed to Class C only | 2
DAS Tools | DIAG_DAS_TOOLS | DAS Tools |  | DataAcquisition |  | X |  |  | X |  |  |  | X |  |  | X |  |  | 1900-01-02 |  |  |  | 4 |  |  | X |  |  | X |  |  | 4
DAS Tools - QIF Vis | DIAG_DAS_TOOLS_QIF_VIS | [7 empty] | X | [13 empty] | 4 | [5 empty] | X |  |  | 4
Scan Data Path | DIAG_SCAN_DATA_PATH | Scan Data Path |  | DataAcquisition |  |  |  | X | [6 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs. |  |  | 2 |  |  |  |  | X |  |  |  | 2
Diagnostic Data Collection (DDC) | DIAG_DATA_COLLECTION | Diagnatoc Data Collection (DDC) |  | DataAcquisition |  |  |  | X | [6 empty] | X |  |  |  | 1900-01-03 | DDC tool has problem with L3 protection, so changed to Level 4. |  |  | 4 |  |  |  |  | X |  |  |  | 4
DIP Diagnostics | DIAG_DIP | DIP Diagnostics |  | DataAcquisition |  |  |  | X | [6 empty] | X |  |  |  | 1900-01-02 |  |  |  | 1 |  |  |  |  | X |  |  |  | 1
High Speed Disk Test | DIAG_HIGH_SPEED_DISK | High Speed Disk Test |  | DataAcquisition |  |  |  |  | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 1 | [5 empty] | X |  |  | 1
Recon Data Path | DIAG_RECON_DATA_PATH | Recon Data Path |  | Image Generation |  |  |  | X | [6 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs. |  |  | 2 |  |  |  |  | X |  |  |  | 2
scsistat OC | DIAG_SCSISTAT_OC | scsistat OC |  | Data Management |  |  |  |  | X | [6 empty] | X |  |  | 1899-12-31 | This command is used by software startup, cannot protect it. |  |  | 1 | [5 empty] | X |  |  | 1
rhapSnap | DIAG_RHAP_SNAP | rhapSnap |  | Data Management |  |  |  |  | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 1 | [5 empty] | X |  |  | 1
GSCB Keyboard Test | DIAG_GSCB_KEYBOARD | SCIM Keyboard Test |  | OperatorIO |  |  |  |  | X | [6 empty] | X |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs. |  |  | 2 | [5 empty] | X |  |  | 2
RTS Viewer(Basic) | DIAG_RTS_VIEWER | [21 empty] | 3 |  |  |  |  | X |  |  |  | 3
RTS Viewer(Advance) | DIAG_RTS_VIEWER_C | [21 empty] | 4 | [5 empty] | X |  |  | 4
RTS Log Viewer |  |  |  |  | X | [17 empty] | na | [8 empty] | 3
System Troubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | System Toubleshooting Tool |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Console Diagnostics | DIAG_CONSOLE | Console Diagnostics | [6 empty] | X | [6 empty] | X |  |  | 1900-01-02 | Need to include DIP Diag in license |  |  | 1 | [5 empty] | X |  |  | 1
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | System Health Monitor | [8 empty] | X | [6 empty] | X | 1900-01-02 | Whether we can wrap /vobs/srv/Tools/shm/UI/shmUI/dist/shmui.jar too? |  |  | 3 | [7 empty] | X | 3
LB Tube Rotor Control | DIAG_LB_TUBE_ROTOR_CONTROL |  |  |  |  | X | [16 empty] | 2 | LB ONLY
LB Tube Rotor Control | DIAG_LB_TUBE_ROTOR_CONTROL_C | [7 empty] | X | [13 empty] | 2 | LB ONLY
LB Tube Rotor Phase Check | [8 empty] | X
Smart KV Tool | DIAG_SMART_KV_TOOL | [28 empty] | X
Image Quality (CSD menu)
Image Analysis | IQ_IMAGE_ANALYSIS | Image Analysis |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | Install SMPTE Image |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Scan Analysis | IQ_SCAN_ANALYSIS | Scan Analysis | [5 empty] | X | [6 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package. |  |  | 2 |  |  |  |  | X |  |  |  | 2
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION | Install Image Orientation Imgs |  |  |  | X | [6 empty] | X | [5 empty] | NA | No protection to the images. |  |  | 3 |  |  | X | [5 empty] | 3
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES |  |  |  | X | [6 empty] | X | [6 empty] | 1900-01-02 | No protection to the images. |  |  | na |  | X | [6 empty] | na
Calibration (CSD menu)
Generator Filament Calibration | CAL_GENERATOR_FILAMENT | Generator Filament Calibration |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 4 |  |  | X | [5 empty] | 4
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION | Gantry Rotation Characterization |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 4 |  |  | X | [5 empty] | 4
Collimator Cal | CAL_COLLIMATOR | Collimator Cal |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-03 |  |  |  | 4 |  |  | X | [5 empty] | 4
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT | Hot ISO Alignment |  |  |  |  | X | [6 empty] | X |  |  |  |  | 1900-01-02 |  |  |  | 4 |  |  |  | X |  |  |  |  | 4
ISO Alignment | CAL_ISO_ALIGNMENT | ISO Alignment |  |  |  |  | X | [6 empty] | X |  |  |  |  | 1900-01-02 |  |  |  | 4 |  |  |  | X |  |  |  |  | 4
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION | Mechanical Characterization |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 4 |  |  | X | [5 empty] | 4
Default Patient Positioning | CAL_DEFAULT_PATIENT_POSITIONING | [23 empty] | X | [6 empty] | na
Detailed Calibration | CAL_DETAILED | Detailed Calibration |  |  |  | X | [7 empty] | X |  |  |  |  | 1900-01-03 |  |  |  | 4 |  |  | X | [5 empty] | 4 | A2
BOW Alignment | CAL_BOW_ALIGNMENT |  |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 4 |  |  | X | [5 empty] | 4
POR Alignment | CAL_POR_ALIGNMENT |  |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 3 |  |  | X | [5 empty] | 3
CBF and SAG Alignment | CAL_CBF_SAG_ALIGNMENT |  |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 4 |  |  | X | [5 empty] | 4
Scan Utilties - Center Phantom | CAL_CENTER_PHANTOM |  |  |  | X | [17 empty] | 4 | on Scan utilities | X | [6 empty] | 4 | on Scan utilities
Z Slope Cal (if no cals loaded) | CAL_Z_SLOPE | [5 empty] | X | [6 empty] | X |  |  |  |  | 1900-01-03 | This is a button on Scan utilities. |  |  | 4 | on Scan utilities |  |  | X |  |  |  |  | 4 | on Scan utilities
DAS Gain Cal | CAL_DAS_GAIN |  |  |  |  | X | [7 empty] | x |  |  |  |  | 1900-01-03 | This is a button on Scan utilities. |  |  | 4 | on Scan utilities |  | X | [5 empty] | 4 | on Scan utilities
P Cal (if no cals loaded) | CAL_PCAL |  |  |  | X | [6 empty] | x | [6 empty] | 1900-01-03 | This is a button on Scan utilities. |  | It is not in the SRS DOC0389957 Rev29. This is shown in Qin software 11BW46.3 | 4 | on Scan utilities | X | [6 empty] | 4 | on Scan utilities
Air Cal (if no cals loaded) | CAL_AIR |  |  |  | X | [6 empty] | x | [6 empty] | 1900-01-03 | This is a button on Scan utilities. |  | It is not in the SRS DOC0389957 Rev29. This is shown in Qin software 11BW46.3 | 4 | on Scan utilities | X | [6 empty] | 4 | on Scan utilities
Scan Utilties - Spectral Cal | CAL_SPECTRAL | [25 empty] | X |  |  |  |  | 4 | on Scan utilities
Configuration (CSD menu)
Config Tracker | CONFIG_TRACKER | Config Tracker |  |  |  | X | [6 empty] | X | [5 empty] | 1899-12-31 | This tool is just a display of config info (which is generated after software LFC). Currently we cannot protect it. |  |  | 1 |  |  | X | [5 empty] | 1
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | Injector: Pressure Limit Unit |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Install Options | CONFIG_INSTALL_OPTIONS | Install Options |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Product SW Revs | CONFIG_PRODUCT_SW_REVS | Product SW Revs |  |  | X | [7 empty] | X | [5 empty] | 1899-12-31 | This command is used by software startup, cannot protect it. |  |  |  | removed
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID | Enter Detector ID |  |  |  | X | [6 empty] | X | [5 empty] | 1900-01-02 |  |  |  | 3 |  |  | X | [5 empty] | 3
Install Camera | CONFIG_INSTALL_CAMERA | Install Camera |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Verify Options | CONFIG_VERIFY_OPTIONS | Verify Options |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Tube Install Certification with Non-GE Tube | CONFIG_TUBE_INSTALL_CERTIFICATION | Tube Install Certification |  |  |  |  | X |  |  |  |  | X | [6 empty] | 1900-01-01 |  |  |  | 3 |  |  |  | X |  |  |  |  | 3
Tube Install Certification wirh GE Tube | CONFIG_TUBE_INSTALL_CERTIFICATION | [9 empty] | X | [11 empty] | 3 | [7 empty] | X | 3
Configure PNF | CONFIG_PNF | Configure PNF |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Configure Enhanced PNF | CONFIG_ENH_PNF |  |  |  | X | [17 empty] | na |  | X | [6 empty] | na
Configure EA3 | CONFIG_EA3 | Configure EA3 |  |  | X | [6 empty] | X | [9 empty] | Not exist in BSD CSD | na |  | X | [6 empty] | na
Configure EAT | CONFIG_EAT | Configure EAT |  |  | X | [6 empty] | X | [9 empty] | Not exist in BSD CSD | na |  | X | [6 empty] | na
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | Configure AW DirectConnect |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Configure AW Server Client | NA |  |  |  | X | [17 empty] | na |  | X | [6 empty] | na
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | Install Gating Monitor |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Component Tracker
Console | CONFIG_COMPONENT_TRACKER_CONSOLE | Console |  | Component Tracker |  |  |  |  | X | [6 empty] | X |  |  | 1899-12-31 |  |  |  | 1 | [5 empty] | X |  |  | 1
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE | Software |  | Component Tracker |  |  |  |  | X | [6 empty] | X |  |  | 1899-12-31 |  |  |  | 1 | [5 empty] | X |  |  | 1
History | CONFIG_COMPONENT_TRACKER_HISTORY | History |  | Component Tracker |  |  |  |  | X | [6 empty] | X |  |  | 1899-12-31 |  |  |  | 1 | [5 empty] | X |  |  | 1
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA | Refresh Console Data |  | Component Tracker |  |  |  |  | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 1 | [5 empty] | X |  |  | 1
InSite Status | CONFIG_INSITE_STATUS |  |  |  | X | [17 empty] | na | add for RSVP | X | [6 empty] | na
Configure InSite | CONFIG_INSITE | [7 empty] | X | [13 empty] | 3 | add for RSVP |  |  |  |  | X |  |  | 3
Configure PRODIAGS | CONFIG_PRODIAGS | [9 empty] | X | [11 empty] | 2 | add for RSVP | [6 empty] | X | 2
Configure Software Download | CONFIG_SW_DOWNLOAD |  |  |  |  | X | [16 empty] | 2 | add for RSVP |  | X | [5 empty] | 2
Configure AWS on VM | None
AWS Client Install/Uninstall | CONFIG_AWVS_CLIENT_INSTALL |  |  |  | X | [19 empty] | X
Configure AWS Applications | CONFIG_AWVS_APPLICATIONS |  |  |  | X | [19 empty] | X
Launch AWS Service Tool | CONFIG_LAUNCH_AWVS_SERVICE |  |  |  | X | [19 empty] | X
Gantry Table User Preference | CONFIG_GANTRY_TABLE_PREFS |  |  |  | X | [17 empty] | na |  | X | [6 empty] | na
Configure HIPAA | CONFIG_HIPAA |  |  |  | X | [17 empty] | na |  | X | [6 empty] | na
Utilities (CSD menu)
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | FLASH Download Tool |  | Install | X | [7 empty] | X | [5 empty] | 1900-01-01 | Technical difficulty to protect the java package. |  |  | na |  | X | [6 empty] | na | A1
HHS Scans | UTIL_HHS_SCANS | HHS Scans |  | Tools |  |  |  |  | X |  |  |  | X |  |  | X |  |  | 1900-01-02 |  |  |  | 3 | [5 empty] | X |  |  | 3
HHS Report | UTIL_HHS_REPORT | HHS Report |  | Tools |  |  |  |  | X |  |  |  | X |  |  | X |  |  | 1900-01-02 |  |  |  | 3 | [5 empty] | X |  |  | 3
Manual CT # Adjust | UTIL_MANUAL_CT_ADJUST |  |  | Tools |  | X | [16 empty] | 4
Auto CT # Adjust | CAL_WATER | Auto CT # Adjust |  | Tools |  | X | [6 empty] | X |  |  | X |  |  | 1900-01-03 |  |  |  | 4 |  |  | X | [5 empty] | 4
Heat Soak and Seasoning | UTIL_HEAT_SOAK_SEASONING | Heat Soak and Seasoning |  | Tools |  | X | [6 empty] | X |  |  | X |  |  | 1900-01-02 |  |  |  | 3 |  |  | X | [5 empty] | 3
CT Mobile | UTIL_CT_MOBILE | CT Mobile |  | Tools | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Verify Security | UTIL_VERIFY_SECURITY | Verify Security |  | Tools | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Editor | UTIL_EDITOR | Editor |  | Utilities |  |  |  | X | [6 empty] | X |  |  |  | 1899-12-31 | Linux built-in or open source tool, no need to protect it. |  |  | 1 |  |  |  |  | X |  |  |  | 1
Calculator | UTIL_CALCULATOR | Calculator |  | Utilities | X | [9 empty] | X |  |  |  | 1899-12-31 | Linux built-in or open source tool, no need to protect it. |  |  |  | removed
Calendar | UTIL_CALENDAR | Calendar |  | Utilities | X | [10 empty] | X |  |  | 1899-12-31 | Linux built-in or open source tool, no need to protect it. |  |  |  | removed | X | [6 empty] | 1
Shell | UTIL_SHELL | Shell |  | Utilities | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Change Display IP | UTIL_CHANGE_DISPLAY_IP | Change Display IP |  | Utilities | X | [9 empty] | X |  |  |  | 1899-12-31 | This function writes the IP to a temp file. No protection to the temp file. No need to protect the web page. It will be update to be identical with CJ later |  |  | na |  | X | [6 empty] | na
Query Updates | UTIL_QUERY_UPDATES | Query Updates |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
Download Updates | UTIL_DOWNLOAD_UPDATES | Download Updates |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
Update Status | UTIL_UPDATE_STATUS | Update Status |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
Update Save | UTIL_UPDATE_SAVE | Update Save |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
Update Restore | UTIL_UPDATE_RESTORE | Update Restore |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
Enable/Disable Auto Install | UTIL_ENABLE_DISABLE_AUTO_INSTALL | Enable/Disble Auto Install |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
Auto Download/Install Status | UTIL_AUTO_DOWNLOAD_INSTALL_STATUS | Auto Download/Install Status |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
Enable/Disble Auto Download | UTIL_ENABLE_DISABLE_AUTO_DOWNLOAD | Enable/Disble Auto Download |  | Update Management |  |  |  | X | [7 empty] | X |  |  | 1900-01-02 |  |  |  |  | removed
System State | UTIL_SYSTEM_STATE | System State |  |  | X | [6 empty] | X | [11 empty] | removed
System State - USB | UTIL_SYSTEM_STATE_USB |  |  |  | X | [6 empty] | X | [10 empty] | na | [8 empty] | na
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | HIPAA setup utility, setting up user accounts |  |  | X | [6 empty] | X | [10 empty] | na | [8 empty] | na
TVA | UTIL_TVA |  |  |  | X | [17 empty] | na | add for Z8G4 | X | [6 empty] | na
Install Software Patch | UTIL_INSTALL_SW_PATCH | [9 empty] | X | [11 empty] | 1 | add for Z8G5 | [6 empty] | X | 1
Security Center | UTIL_SECURITY_CENTER |  |  |  | X | [17 empty] | na | add for Z8G6 | X | [6 empty] | na
Reset Admin User Password | UTIL_RST_ADM_PASSWD | [9 empty] | X | [11 empty] | na | [7 empty] | X |  | Added by HCSDM00633801.
Reset Service User Password | UTIL_RST_SRV_PASSWD | [9 empty] | X | [11 empty] | na | [7 empty] | X |  | Added by HCSDM00633801.
Audit Log Viewer |  |  |  |  | X | [17 empty] | na |  | X | [6 empty] | na
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | Application Shutdown |  |  | X | [6 empty] | X | [10 empty] | na |  | X | [6 empty] | na
Replacement (CSD menu)
New Tube Config | REP_NEW_TUBE_CONFIG | New Tube Config |  |  |  |  | X | [6 empty] | X |  |  |  |  | 1900-01-02 |  |  |  | 3 |  |  |  | X |  |  |  |  | 3
Product Locator |  |  |  |  | X | [17 empty] | na |  | X | [6 empty] | na
PM (CSD menu)
Service Notepad | PM_SERVICE_NOTEPAD | Service Notepad | [6 empty] | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 3 | [5 empty] | X |  |  | 3
Session Log | PM_SESSION_LOG | Session Log | [6 empty] | X | [6 empty] | X |  |  | 1900-01-01 | Technical difficulty to protect the java package. |  |  | 2 | [5 empty] | X |  |  | 2
System Health Page | PM_SYSTEM_HEALTH_PAGE | System Health Page | [6 empty] | X | [6 empty] | X |  |  | 1900-01-02 |  |  |  | 3 | [5 empty] | X |  |  | 3
Note :
if a tool in both ClassA and ClassC, that means it will appear in both scenarios
Upper folder can only appear when system is ClassC or ClassM licensed
This CSD-Service Tool List covers the product of BSD(DOC0928984 BSD Series Production History and IB Configuration Summary_Rev6) and KL-PDAS
Note on E2: if a tool has upper folder, that means it will appear under the upper folder when this folder comes out. Folder comes our if it is classC/M licensed
Note on A27: Tao, Zhou (GE Healthcare): Need classification on this item. This tool was enabled in DAS tools for Class C.
Note on W28: class A runs the test and give errors in syslog==>give away a lot -
Note on A30: One of the most powerfull tools. For engineeirng protocle name, tracking…you can make it do the coolimeter test, the axial test.not explained in process - should be in class M.there is no procedure for it.
Note on W30: One of the most powerfull tools. For engineeirng protocle name, tracking…you can make it do the coolimeter test, the axial test.not explained in process - should be in class M.there is no procedure for it.

## KL16 Tools&Utilities List
KL16 Service tools reclassification
view Service Document |  |  |  | Current IB Classification |  |  |  | New IB Classification - Cj
System tools/utilities/diags (As they exist on KL16 today, very similar to older CT IB systems as well) | Tool License Token | CJ tools name | Tool overview |  | Class A | Class C | Class M |  | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | Protection Level | Comments
Error Logs (CSD menu)
System Error Log | ERROR_SYSTEM_LOG | System Error Log |  |  |  | X | [6 empty] | X |  |  | 2 | No protection to log files anyway. And it's using a viewlog command which may be widely used.
Software History Logs | ERROR_SOFTWARE_HISTORY_LOGS | Software History Log |  |  |  | X | [6 empty] | X |  |  | 3
Store Log | ERROR_STORE_LOG | Storelog |  |  | X |  |  |  |  | X | [5 empty] | 3
System Browser(All feature) | ERROR_SYSTEM_BROWSER | System Browser (All features ) |  |  | X | [6 empty] | X |  |  |  | 1 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
Tube Usage | ERROR_TUBE_USAGE | Tube Usage |  |  | X | [6 empty] | X |  |  |  | 1 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
LogViewer | ERROR_LOG_VIEWER | LogViewer |  |  | X | [6 empty] | X |  |  |  | 3
StartGFV | ERROR_START_GFV | StartGFV |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty: it's only a tar package installed to console.
Diagnostics (CSD menu)
LAN Diagnositics | DIAG_LAN | LAN Diagnostics |  |  |  | X | [6 empty] | X |  |  | 3
CPU Processes | DIAG_CPU_PROCESSES | CPU Processes |  |  |  | X | [6 empty] | X |  |  | 1 | This is a Linux built-in command. No need to protect it.
kV & mA (X-ray) | DIAG_KV_MA | kV & mA (X-ray) |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI | Generator Tool - JEDI |  |  | X | X | X |  |  | X |  |  | X |  | X | 3 | Need to include KVMA in license
Generator Tool JEDI - Focal Spot | DIAG_FOCAL_SPOT | [14 empty] | 4 | Need classification.
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE | Collimator Aperture Test |  |  | X | [6 empty] | X |  |  |  | 3
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL | Collimator Functional |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Axial Functional | DIAG_AXIAL_FUNCTIONAL | Axial Functional |  |  |  | X | [6 empty] | X |  |  | 3
DAS Tools | DIAG_DAS_TOOLS | DAS Tools |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
DAS Tools - QIF Vis | DIAG_DAS_TOOLS_QIF_VIS |  |  |  |  | X | [9 empty] | 4 | Need classification.
Scan Data Path | DIAG_SCAN_DATA_PATH | Scan Data Path |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Diagnostic Data Collection (DDC) | DIAG_DATA_COLLECTION | Diagnatoc Data Collection (DDC) |  |  | X | [6 empty] | X |  |  |  | 4 | DDC tool has problem with L3 protection, so changed to Level 4.
DIP Diagnostics | DIAG_DIP | VDIP Diagnostics |  |  | X | [6 empty] | X |  |  |  | 3
High Speed Disk Test | DIAG_HIGH_SPEED_DISK | High Speed Disk Test |  |  |  | X | [6 empty] | X |  |  | 3
Recon Data Path | DIAG_RECON_DATA_PATH | Recon Data Path |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
RAC Diagnostics | DIAG_RAC | VRAC Diagnostics | [13 empty] | 3
scsistat OC | DIAG_SCSISTAT_OC | scsistat OC |  |  |  | X | [6 empty] | X |  |  | 1 | This command is used by software startup, cannot protect it.
rhapSnap | DIAG_RHAP_SNAP | rhapSnap |  |  |  | X | [6 empty] | X |  |  | 3
GSCB Keyboard Test | DIAG_GSCB_KEYBOARD | SCIM Keyboard Test |  |  |  | X | [6 empty] | X |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
RTS viewer | DIAG_RTS_VIEWER | RTS Viewer A |  |  | X | [6 empty] | X |  |  |  | 3
RTS viewer C | DIAG_RTS_VIEWER_C | RTS Viewer C |  |  |  | X | [6 empty] | X |  |  | 4
System Troubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | System Toubleshooting Tool |  |  | X |  |  |  | X
Console Diagnostics | DIAG_CONSOLE | Console Diagnostics |  |  |  | X | [6 empty] | X |  |  | 3 | Need to include DIP Diag in license
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | System Health Monitor |  |  |  |  | X | [7 empty] | X | 3 | Whether we can wrap /vobs/srv/Tools/shm/UI/shmUI/dist/shmui.jar too?
Smart kV Tool | DIAG_SMART_KV_TOOL | Smart kV Tool |  |  |  |  | X | [6 empty] | X |  | 3
Image Quality (CSD menu)
Image Analysis | IQ_IMAGE_ANALYSIS | Image Analysis |  |  | X |  |  |  | X
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | Install SMPTE Image |  |  | X |  |  |  | X
Scan Analysis | IQ_SCAN_ANALYSIS | Scan Analysis |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package.
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION | Install Image Orientation Imgs |  |  | X |  |  |  |  | X | [5 empty] | NA | No protection to the images.
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES |  |  |  | X |  |  |  | X | [6 empty] | 3 | No protection to the images.
Calibration (CSD menu)
Generator Filament Calibration | CAL_GENERATOR_FILAMENT | Generator Filament Calibration |  |  | X |  |  |  |  | X | [5 empty] | 3
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION | Gantry Rotation Characterization |  |  | X |  |  |  |  | X | [5 empty] | 3
Collimator Cal | CAL_COLLIMATOR | Collimator Cal |  |  | X |  |  |  |  | X | [5 empty] | 4
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT | Hot ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 3
ISO Alignment | CAL_ISO_ALIGNMENT | ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 3
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION | Mechanical Characterization |  |  | X |  |  |  |  | X | [5 empty] | 3
Default Patient Positioning | CAL_DEFAULT_PATIENT_POSITIONING | Default Patient Positioning |  |  | X |  |  |  | X
Detailed Calibration | CAL_DETAILED | Detailed Calibration |  |  | X | [5 empty] | X |  |  |  |  | 4
BOW Alignment | CAL_BOW_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
POR Alignment | CAL_POR_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
CBF and SAG Alignment | CAL_CBF_SAG_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
| CAL_Z_ALIGN | Z Align |  |  | X |  |  |  |  | X | [5 empty] | TBD | Tool not used by KL16, TBD for CJ.
| CAL_SMART_TUBE_ALIGNMENT | Smart Tube Alignment |  |  |  |  | X | [6 empty] | X |  | TBD | Tool not used by KL16, TBD for CJ.
| CAL_GANTRY_SERVICE_BALANCE | Gantry Service Balance |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package.
Scan Utilties - Center Phantom | CAL_CENTER_PHANTOM |  |  |  | x |  |  |  | x | [6 empty] | 4 | This is a button on Scan utilities.
Scan Utilties - Z Slope Cal | CAL_Z_SLOPE |  |  |  | x | [5 empty] | X |  |  |  |  | 4 | This is a button on Scan utilities.
Scan Utilties - DAS Gain Cal | CAL_DAS_GAIN |  |  |  | x | [5 empty] | x |  |  |  |  | 4 | This is a button on Scan utilities.
Scan Utilties - Air Cal | CAL_AIR |  |  |  | x |  |  |  | x | [6 empty] | 4 | This is a button on Scan utilities.
Scan Utilties - Spectral Cal | CAL_SPECTRAL |  |  |  | x | [5 empty] | x |  |  |  |  | 4 | This is a button on Scan utilities.
Configuration (CSD menu) |  | Transfer Function Calibration
Config Tracker | CONFIG_TRACKER | Config Tracker |  |  | X |  |  |  |  | X | [5 empty] | 1 | This tool is just a display of config info (which is generated after software LFC). Currently we cannot protect it.
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | Injector: Pressure Limit Unit |  |  | X |  |  |  | X
Install Options | CONFIG_INSTALL_OPTIONS | Install Options |  |  | X |  |  |  | X
Product SW Revs | CONFIG_PRODUCT_SW_REVS | Product SW Revs |  |  | X |  |  |  |  | X | [5 empty] | 1 | This command is used by software startup, cannot protect it.
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID | Enter Detector ID |  |  | X |  |  |  |  | X | [5 empty] | 3
Install Camera | CONFIG_INSTALL_CAMERA | Install Camera |  |  | X |  |  |  | X
Verify Options | CONFIG_VERIFY_OPTIONS | Verify Options |  |  | X |  |  |  | X
Tube Install Certification | CONFIG_TUBE_INSTALL_CERTIFICATION | Tube Install Certification |  |  | X |  | X |  |  |  | X |  |  |  | X | 2
Configure PNF | CONFIG_PNF | Configure PNF |  |  | X |  |  |  | X
Configure EA3 | CONFIG_EA3 | Configure EA3 |  |  | X |  |  |  | X
Configure EAT | CONFIG_EAT | Configure EAT |  |  | X |  |  |  | X
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | Configure AW DirectConnect |  |  | X |  |  |  | X
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | Install Gating Monitor |  |  | X |  |  |  | X
Component Tracker |  | Component Tracker
Console | CONFIG_COMPONENT_TRACKER_CONSOLE | Console |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE | Software |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
History | CONFIG_COMPONENT_TRACKER_HISTORY | History |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA | Refresh Console Data |  |  |  | X | [6 empty] | X |  |  | 3
| CONFIG_COMPONENT_TRACKER_DETECTOR | Detector | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
|  | Configure AW Server Client | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
|  | Patient Information on Xtream Display | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
Utilities (CSD menu)
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | FLASH Download Tool |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package.
HHS Scans | UTIL_HHS_SCANS | HHS Scans |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
HHS Report | UTIL_HHS_REPORT | HHS Report |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
Auto CT # Adjust | UTIL_AUTO_CT_NUMBER_ADJUST | Auto CT # Adjust |  |  | X | X |  |  |  | X |  |  | X |  |  | 4
Heat Soak and Seasoning | UTIL_HEAT_SOAK_SEASONING | Heat Soak and Seasoning |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
CT Mobile | UTIL_CT_MOBILE | CT Mobile |  |  | X |  |  |  | X
Verify Security | UTIL_VERIFY_SECURITY | Verify Security |  |  | X |  |  |  | X
Editor | UTIL_EDITOR | Editor |  |  | X | [6 empty] | X |  |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Calculator | UTIL_CALCULATOR | Calculator |  |  | X | [6 empty] | X |  |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Calendar | UTIL_CALENDAR | Calendar |  |  |  | X | [6 empty] | X |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Shell | UTIL_SHELL | Shell |  |  | X |  |  |  | X
Change Display IP | UTIL_CHANGE_DISPLAY_IP | Change Display IP |  |  | X | [6 empty] | X |  |  |  | 1 | This function writes the IP to a temp file. No protection to the temp file. No need to protect the web page. It will be update to be identical with CJ later
Query Updates | UTIL_QUERY_UPDATES | Query Updates |  |  |  | X | [6 empty] | X |  |  | 3
Download Updates | UTIL_DOWNLOAD_UPDATES | Download Updates |  |  |  | X | [6 empty] | X |  |  | 3
Update Status | UTIL_UPDATE_STATUS | Update Status |  |  |  | X | [6 empty] | X |  |  | 3
Update Save | UTIL_UPDATE_SAVE | Update Save |  |  |  | x | [6 empty] | X |  |  | 3
Update Restore | UTIL_UPDATE_RESTORE | Update Restore |  |  |  | X | [6 empty] | X |  |  | 3
Enable/Disable Auto Install | UTIL_ENABLE_DISABLE_AUTO_INSTALL | Enable/Disble Auto Install |  |  |  | X | [6 empty] | X |  |  | 3
Auto Download/Install Status | UTIL_AUTO_DOWNLOAD_INSTALL_STATUS | Auto Download/Install Status |  |  |  | X | [6 empty] | X |  |  | 3
Enable/Disble Auto Download | UTIL_ENABLE_DISABLE_AUTO_DOWNLOAD | Enable/Disble Auto Download |  |  |  | X | [6 empty] | X |  |  | 3
System State | UTIL_SYSTEM_STATE | System State |  |  | X |  |  |  | X
System State - USB | UTIL_SYSTEM_STATE_USB |  |  |  | X |  |  |  | X
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | EA3 Admin Browser |  |  | X |  |  |  | X
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | Application Shutdown |  |  | X |  |  |  | X
Replacement (CSD menu)
New Tube Config | REP_NEW_TUBE_CONFIG | New Tube Config |  |  | X | [5 empty] | X |  |  |  |  | 3
FRDM Wizard | REP_FRDM_WIZARD | FRDM Wizard |  |  | X |  |  |  |  | X | [5 empty] | 1 | This tool is not used for KL16 China. Change level to 1 for CJ.
PM (CSD menu)
Service Notepad | PM_SERVICE_NOTEPAD | Service Notepad |  |  |  | X | [6 empty] | X |  |  | 3
Session Log | PM_SESSION_LOG | Session Log |  |  |  | X | [6 empty] | X |  |  | 2 | Technical difficulty to protect the java package.
System Health Page | PM_SYSTEM_HEALTH_PAGE | System Health Page |  |  |  | X | [6 empty] | X |  |  | 3
Note on A14: Tao, Zhou (GE Healthcare): A Linux built-in command, no need to protect
Note on C16: 212007413: Check A,C, and M functionality. Question if there are separate functions for each class
Note on A17: Tao, Zhou (GE Healthcare): A tool on Generator Tool JEDI which is enabled for Class C leve. Now used a special token
Note on C18: check if required in replacement or install procedure take a serie of 8/10 scans
Note on C19: eng use it.used for troubleshooting.
Note on C21: develop a reduce set of class A QIF only for Class C
Note on A22: Tao, Zhou (GE Healthcare): Need classification on this item. This tool was enabled in DAS tools for Class C.
Note on U23: class A runs the test and give errors in syslog==>give away a lot -
Note on U25: One of the most powerfull tools. For engineeirng protocle name, tracking…you can make it do the coolimeter test, the axial test.not explained in process - should be in class M.there is no procedure for it.
Note on A28: 212007413: Check to see if Kunlun16 with NIO has RAC diags. Check RT16
Note on C34: Dave: should it be moved to C ? As by replacing the TGP…or using a miror to see the LEDs on the TGP board. ( table gantry processor)
Note on U37: phantom need to be relabeleed, sw change with sw doc, service procedure & reference to procedure sw( 2 sw : 1 matlab and 1 accs version, old C code)+doc = 1man month depends on sw release
Note on C52: all calibration
Note on C57: Mkg would like to move it to class C
Note on U58: Filliat, Claire (GE Healthcare): need added work ot move to C : sw and service doc
Note on C64: required for tube + tank
Note on A90: Tao, Zhou (GE Healthcare): No A level for HHS Scans and HHS Report?
Note on A115: Tao, Zhou (GE Healthcare): Tool not found on KL16

## BSD-Z840 CSD Menu
BSD Service tools | [15 empty] | #NAME?
view Service Document |  |  |  | Current IB Classification |  |  |  | New IB Classification - Cj
System tools/utilities/diags | Tool License Token | Tools name | Tool overview | Upper Folder | Class A | Class C | Class M |  | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | Protection Level | Marks | Impact of reclassification to Indepedant Service Provider (Class A moved to Class C) | Difference | Comments
Error Logs (CSD menu)
System Error Log | ERROR_SYSTEM_LOG | System Error Log |  |  |  | X | [6 empty] | X |  |  | 1900-01-01 | No protection to log files anyway. And it's using a viewlog command which may be widely used.
Software History Logs | ERROR_SOFTWARE_HISTORY_LOGS | Software History Log |  |  |  | X | [6 empty] | X |  |  | 1900-01-02
Store Log | ERROR_STORE_LOG | Storelog |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
System Browser(All feature) | ERROR_SYSTEM_BROWSER | System Browser (All features ) |  |  | X | [6 empty] | X |  |  |  | 1899-12-31 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
Tube Usage | ERROR_TUBE_USAGE | Tube Usage |  |  | X | [6 empty] | X |  |  |  | 1899-12-31 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
LogViewer | ERROR_LOG_VIEWER | LogViewer |  |  | X | [6 empty] | X |  |  |  | 1900-01-02
StartGFV | ERROR_START_GFV | StartGFV |  |  |  | X | [6 empty] | X |  |  | 1899-12-31 | Technical difficulty: it's only a tar package installed to console.
Diagnostics (CSD menu)
LAN Diagnositics (System Communication) | DIAG_LAN | LAN Diagnostics |  | SystemCommunications |  | X | [6 empty] | X |  |  | 1900-01-02
CPU Processes (System Monitoring) | DIAG_CPU_PROCESSES | CPU Processes |  | SystemMonitoring |  | X | [6 empty] | X |  |  | 1899-12-31 | This is a Linux built-in command. No need to protect it.
kV & mA (X-ray Generation) | DIAG_KV_MA | kV & mA (X-ray) |  | X-RayGeneration | X | X |  |  |  | X | [5 empty] | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs.
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI | Generator Tool - JEDI |  | X-RayGeneration | X | X | X |  |  | X |  |  | X |  | X | 1900-01-02 | Need to include KVMA in license
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE | Collimator Aperture Test |  | X-RayGeneration | X | [6 empty] | X |  |  |  | 1900-01-02
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL | Collimator Functional |  | X-RayGeneration | X | X | [5 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs.
Axial Functional (AxialControl) | DIAG_AXIAL_FUNCTIONAL | Axial Functional |  | AxialControl |  | X | [6 empty] | X |  |  | 1900-01-02
Gantry Service Balance (AxialControl) |  |  |  | AxialControl |  | X | [12 empty] | Per SWRS.CSD.250, this was once defined as both Class A/C, don't know why later changed to Class C only | Suggest to let it be ClassC
DAS Tools | DIAG_DAS_TOOLS | DAS Tools |  | DataAcquisition | X | X |  |  |  | X |  |  | X |  |  | 1900-01-02
Scan Data Path | DIAG_SCAN_DATA_PATH | Scan Data Path |  | DataAcquisition | X | X | [5 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs.
Diagnostic Data Collection (DDC) | DIAG_DATA_COLLECTION | Diagnatoc Data Collection (DDC) |  | DataAcquisition | X | X | [5 empty] | X |  |  |  | 1900-01-03 | DDC tool has problem with L3 protection, so changed to Level 4.
DIP Diagnostics | DIAG_DIP | DIP Diagnostics |  | DataAcquisition | X | X | [5 empty] | X |  |  |  | 1900-01-02
High Speed Disk Test | DIAG_HIGH_SPEED_DISK | High Speed Disk Test |  | DataAcquisition |  | X | [6 empty] | X |  |  | 1900-01-02
Recon Data Path | DIAG_RECON_DATA_PATH | Recon Data Path |  | Image Generation | X | X | [5 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs.
RAC Diagnostics | DIAG_RAC | RAC Diagnostics |  | Image Generation |  | X | [9 empty] | 1900-01-02
scsistat OC | DIAG_SCSISTAT_OC | scsistat OC |  | Data Management |  | X | [6 empty] | X |  |  | 1899-12-31 | This command is used by software startup, cannot protect it.
rhapSnap | DIAG_RHAP_SNAP | rhapSnap |  | Data Management |  | X | [6 empty] | X |  |  | 1900-01-02
GSCB Keyboard Test | DIAG_GSCB_KEYBOARD | SCIM Keyboard Test |  | OperatorIO |  | X | [6 empty] | X |  |  | 1900-01-01 | Technical difficulty to protect the java package because it's shared by many programs. |  |  | The original is called SCIM, suggest to be GSCB
System Troubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | System Toubleshooting Tool |  |  | X | X |  |  | X
Console Diagnostics | DIAG_CONSOLE | Console Diagnostics |  |  |  | X | [6 empty] | X |  |  | 1900-01-02 | Need to include DIP Diag in license
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | System Health Monitor |  |  |  |  | X | [7 empty] | X | 1900-01-02 | Whether we can wrap /vobs/srv/Tools/shm/UI/shmUI/dist/shmui.jar too?
LB Tube Rotor Control |  |  |  |  | X | [14 empty] | For systems equipped with Liquid Bearing Tube ONLY
LB Tube Rotor Phase Check | [6 empty] | X | [12 empty] | For systems equipped with Liquid Bearing Tube ONLY
Image Quality (CSD menu)
Image Analysis | IQ_IMAGE_ANALYSIS | Image Analysis |  |  | X |  |  |  | X
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | Install SMPTE Image |  |  | X |  |  |  | X
Scan Analysis | IQ_SCAN_ANALYSIS | Scan Analysis |  |  | X | [6 empty] | X |  |  |  | 1900-01-01 | Technical difficulty to protect the java package.
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION | Install Image Orientation Imgs |  |  | X |  |  |  |  | X | [5 empty] | NA | No protection to the images.
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES |  |  |  | X |  |  |  | X | [6 empty] | 1900-01-02 | No protection to the images.
Calibration (CSD menu)
Generator Filament Calibration | CAL_GENERATOR_FILAMENT | Generator Filament Calibration |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION | Gantry Rotation Characterization |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
Collimator Cal | CAL_COLLIMATOR | Collimator Cal |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-03
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT | Hot ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 1900-01-02
ISO Alignment | CAL_ISO_ALIGNMENT | ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 1900-01-02
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION | Mechanical Characterization |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
Detailed Calibration | CAL_DETAILED | Detailed Calibration |  |  | X | [5 empty] | X |  |  |  |  | 1900-01-03
BOW Alignment | CAL_BOW_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
POR Alignment | CAL_POR_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
CBF and SAG Alignment | CAL_CBF_SAG_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
Ganry Service Balance | CAL_GANTRY_SERVICE_BALANCE | Gantry Service Balance |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-01 | Technical difficulty to protect the java package.
Z Slope Cal (if no cals loaded) | CAL_Z_SLOPE |  |  |  | x | [5 empty] | X |  |  |  |  | 1900-01-03 | This is a button on Scan utilities.
DAS Gain Cal | CAL_DAS_GAIN |  |  |  | x | [5 empty] | x |  |  |  |  | 1900-01-03 | This is a button on Scan utilities.
P Cal (if no cals loaded) | CAL_AIR |  |  |  | X |  |  |  | x | [6 empty] | 1900-01-03 | This is a button on Scan utilities. |  | It is not in the SRS DOC0389957 Rev29. This is shown in Qin software 11BW46.3 | May need to change to Class A
Air Cal (if no cals loaded) | CAL_AIR |  |  |  | X |  |  |  | x | [6 empty] | 1900-01-03 | This is a button on Scan utilities. |  | It is not in the SRS DOC0389957 Rev29. This is shown in Qin software 11BW46.3 | May need to change to Class A
Configuration (CSD menu)
Config Tracker | CONFIG_TRACKER | Config Tracker |  |  | X |  |  |  |  | X | [5 empty] | 1899-12-31 | This tool is just a display of config info (which is generated after software LFC). Currently we cannot protect it.
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | Injector: Pressure Limit Unit |  |  | X |  |  |  | X
Install Options | CONFIG_INSTALL_OPTIONS | Install Options |  |  | X |  |  |  | X
Product SW Revs | CONFIG_PRODUCT_SW_REVS | Product SW Revs |  |  | X |  |  |  |  | X | [5 empty] | 1899-12-31 | This command is used by software startup, cannot protect it.
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID | Enter Detector ID |  |  | X |  |  |  |  | X | [5 empty] | 1900-01-02
Install Camera | CONFIG_INSTALL_CAMERA | Install Camera |  |  | X |  |  |  | X
Verify Options | CONFIG_VERIFY_OPTIONS | Verify Options |  |  | X |  |  |  | X
Tube Install Certification | CONFIG_TUBE_INSTALL_CERTIFICATION | Tube Install Certification |  |  | X | [5 empty] | X |  |  |  | X | 1900-01-01
Configure PNF | CONFIG_PNF | Configure PNF |  |  | X |  |  |  | X
Configure EA3 | CONFIG_EA3 | Configure EA3 |  |  | X |  |  |  | X | [9 empty] | Not exist in BSD CSD | We need this feature for customer
Configure EAT | CONFIG_EAT | Configure EAT |  |  | X |  |  |  | X | [9 empty] | Not exist in BSD CSD | We need this feature for both Service and customer
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | Configure AW DirectConnect |  |  | X |  |  |  | X
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | Install Gating Monitor |  |  | X |  |  |  | X
Console | CONFIG_COMPONENT_TRACKER_CONSOLE | Console |  | Component Tracker |  | X | [6 empty] | X |  |  | 1899-12-31
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE | Software |  | Component Tracker |  | X | [6 empty] | X |  |  | 1899-12-31
History | CONFIG_COMPONENT_TRACKER_HISTORY | History |  | Component Tracker |  | X | [6 empty] | X |  |  | 1899-12-31
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA | Refresh Console Data |  | Component Tracker |  | X | [6 empty] | X |  |  | 1900-01-02
Configure HIPAA | CONFIG_HIPAA |  |  |  | X | [14 empty] | This needs to add since NEMA XR-26
Utilities (CSD menu)
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | FLASH Download Tool |  | Install | X |  |  |  |  | X | [5 empty] | 1900-01-01 | Technical difficulty to protect the java package.
HHS Scans | UTIL_HHS_SCANS | HHS Scans |  | Tools |  | X |  |  |  | X |  |  | X |  |  | 1900-01-02
HHS Report | UTIL_HHS_REPORT | HHS Report |  | Tools |  | X |  |  |  | X |  |  | X |  |  | 1900-01-02
Manual CT # Adjust |  |  |  | Tools |  | X
Auto CT # Adjust | UTIL_AUTO_CT_NUMBER_ADJUST | Auto CT # Adjust |  | Tools |  | X |  |  |  | X |  |  | X |  |  | 1900-01-03
Heat Soak and Seasoning | UTIL_HEAT_SOAK_SEASONING | Heat Soak and Seasoning |  | Tools |  | X |  |  |  | X |  |  | X |  |  | 1900-01-02
CT Mobile | UTIL_CT_MOBILE | CT Mobile |  | Tools | X |  |  |  | X
Verify Security | UTIL_VERIFY_SECURITY | Verify Security |  | Tools | X |  |  |  | X
Editor | UTIL_EDITOR | Editor |  | Utilities | X | [6 empty] | X |  |  |  | 1899-12-31 | Linux built-in or open source tool, no need to protect it.
Calculator | UTIL_CALCULATOR | Calculator |  | Utilities | X | [6 empty] | X |  |  |  | 1899-12-31 | Linux built-in or open source tool, no need to protect it.
Calendar | UTIL_CALENDAR | Calendar |  | Utilities | X | [7 empty] | X |  |  | 1899-12-31 | Linux built-in or open source tool, no need to protect it.
Shell | UTIL_SHELL | Shell |  | Utilities | X |  |  |  | X
Change Display IP | UTIL_CHANGE_DISPLAY_IP | Change Display IP |  | Utilities | X | [6 empty] | X |  |  |  | 1899-12-31 | This function writes the IP to a temp file. No protection to the temp file. No need to protect the web page. It will be update to be identical with CJ later
Query Updates | UTIL_QUERY_UPDATES | Query Updates |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
Download Updates | UTIL_DOWNLOAD_UPDATES | Download Updates |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
Update Status | UTIL_UPDATE_STATUS | Update Status |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
Update Save | UTIL_UPDATE_SAVE | Update Save |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
Update Restore | UTIL_UPDATE_RESTORE | Update Restore |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
Enable/Disable Auto Install | UTIL_ENABLE_DISABLE_AUTO_INSTALL | Enable/Disble Auto Install |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
Auto Download/Install Status | UTIL_AUTO_DOWNLOAD_INSTALL_STATUS | Auto Download/Install Status |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
Enable/Disble Auto Download | UTIL_ENABLE_DISABLE_AUTO_DOWNLOAD | Enable/Disble Auto Download |  | Update Management |  | X | [6 empty] | X |  |  | 1900-01-02
System State | UTIL_SYSTEM_STATE | System State |  |  | X |  |  |  | X
System State - USB | UTIL_SYSTEM_STATE_USB |  |  |  | X |  |  |  | X
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | HIPAA setup utility, setting up user accounts |  |  | X |  |  |  | X
Audit Log Viewer |  |  |  |  | X | [14 empty] | This needs to add since NEMA XR-26
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | Application Shutdown |  |  | X |  |  |  | X
Replacement (CSD menu)
New Tube Config | REP_NEW_TUBE_CONFIG | New Tube Config |  |  | X | [5 empty] | X |  |  |  |  | 1900-01-02
Product Locator |  |  |  |  | X | [14 empty] | Add this to be identical to other products
PM (CSD menu)
Service Notepad | PM_SERVICE_NOTEPAD | Service Notepad |  |  |  | X | [6 empty] | X |  |  | 1900-01-02
Session Log | PM_SESSION_LOG | Session Log |  |  |  | X | [6 empty] | X |  |  | 1900-01-01 | Technical difficulty to protect the java package.
System Health Page | PM_SYSTEM_HEALTH_PAGE | System Health Page |  |  |  | X | [6 empty] | X |  |  | 1900-01-02
Note :
if a tool in both ClassA and ClassC, that means it will appear in both scenarios
Upper folder can only appear when system is ClassC or ClassM licensed
Note on E3: if a tool has upper folder, that means it will appear under the upper folder when this folder comes out. Folder comes our if it is classC/M licensed
Note on U22: class A runs the test and give errors in syslog==>give away a lot -
Note on A24: One of the most powerfull tools. For engineeirng protocle name, tracking…you can make it do the coolimeter test, the axial test.not explained in process - should be in class M.there is no procedure for it.
Note on U24: One of the most powerfull tools. For engineeirng protocle name, tracking…you can make it do the coolimeter test, the axial test.not explained in process - should be in class M.there is no procedure for it.
Note on A27: 212007413: Check to see if Kunlun16 with NIO has RAC diags. Check RT16

## Revision History
This document is applied to both BrightSpeed Power and Lite systems, including the systems equipped with Liquid Bearing Tube too
Delight LiteSW DHF index
| Rev | Version | Date | Author | Reason for Change/Contens of Change
| 1 | 1 | 2017-01-20 | Li Jiang | Initial Release based on DOC0389957 Rev29
| 1 | 2 | 2017-02-15 | Li Jiang | 1. Deleted Z-align since it is not in new software per checking in bj_sles_cl77 2. Move P-cal and Air-CAL in Calibration Tab from Class C to Class C per new software. 3. Added Product Locator in Replacement Tab as new tool
| 2 | 1 | 2020-09-15 | Wen jianwei | 1.BSD16+KL16-Z8G4 CSD-service tool list initial release with adding sheet -BSD-Z8G4 CSD Menu
| 2 | 1 | 2020-12-28 | Wen jianwei | 1.add RTS Log Viewer-A；remove RTS Viewer 2.add nanoAWS related CSD menu 3.add class C token for LB Tube Rotor Control 4.split the Tube Install Certification to "with Non-GE Tube" and "wirh GE Tube" 5.add token name for Gantry Service Balance for BSD，not for KL 6.add CSD menu “ComponentTtracker” 7.add Reset Admin User Password and Reset Service User Password
| 3 | 1 | 2021-05-18 | Wen jianwei | 1.recover RTS Viewer 2.remove Configure PNF , add Configure Enhanced PNF 3，update the token for Manual CT # Adjust and Auto CT # Adjust
| 4 | 1 | 2021-10-14 | wen jianwei | 1, Remove the tool Calendar 2, Remove the tool Configure AW Server Client