## Sheet1
EC# | Description | Location
EC1 | The system is designed to provide adequate protection against hazards caused by overflow | 1. Service Methods (5366639-8EN) -> Replacement-> Gantry -> Tube -> X-Ray Tube Replacement 2. Service Methods (5366639-8EN) -> Align, Setup, Cals -> Gantry -> Tilt -> Hydraulic Tilt Motor Fluid Check and Fill Procedure
EC2 | The system is designed to provide adequate protection against hazards caused by leakage | Service methods (5366639-8EN)->Functional Checks-> Enclosure Leakage and system chassis leakage tests
EC3 | The system is designed to provide adequate protection against hazards caused by humidity | Pre-Installation Manual 5366636-1EN -> Chapter 4, Section 1.0 Installation Manual 5366637-1EN -> Chapter 1, Section 3.0
EC4 | Covers are designed to minimise the possibility of fluid ingress. | Pre-Installation Manual 5366636-1EN -> Chapter 3, Section 4.0
EC5 | Comply with applicable standards and regulations pertaining to Radiation Control | Pre-Installation Manual 5366636-1EN -> Chapter 3, Section 7.0
EC6 | Comply with applicable standards and regulations pertaining to Xray Indication Safety | Pre-Installation Manual 5366636-1EN -> Chapter 8, Section 7.2
EC7 | Manufacturing and Service testing procedures for kV, mA, and scan time accuracy. | Service Methods (5366639-8EN) -> Functional Checks -> Gantry -> HV Tank Feedback Resistor Verification Service Methods (5366639-8EN) -> Functional Checks -> Gantry -> Meter Verification
EC8 | Design uses IEC standards as a tool for safety factor definition. Manufacturing and Service procedure gives information on how to balance the gantry. | Service Methods (5366639-8EN) -> Align, Setup, Cals-> Gantry-> Gantry Balance Procedure
EC9 | In the event of failure of Service Process, the clearance to the stationary frame should be verified by a field service engineer before rotation and Service Methods shall provide 'Manually rotate gantry, at least one full revolution, and verify that no rubbing occurs' at Tube replacement section. | Service Methods (5366639-8EN) -> Replacement -> Gantry -> Tube -> Tube Replacement
EC10 | The system tolerates the use of biologically effective cleaning agents Cleaning procedure in user and service documentation. | Service Methods (5366639-8EN) -> Planned Maintenance and Tasks
EC11 | Service procedure instructs proper Date and Time setting in software installation process. | Service Methods (5366639-8EN) -> Software -> Load From Cold Procedure
EC12 | Include periodic maintenance, daily and periodic operator checks and calibrations in service and operator manual. Objective Evidence: Service Manuals and User Instructions | Service Methods (5366639-8EN) -> Planned Maintenance and Tasks
EC13 | System provides a PM Schedules document along with the scanner. | Service Methods (5366639-8EN) -> Planned Maintenance and Tasks
EC14 | Service procedure to characterize table/gantry motion at install | Installation Manual 5366637-1EN -> Chapter 5, Section 4.0