## Sheet1
Service tools reclassification |  |  |  | Classification
System tools/utilities/diags | Tool License Token |  | Tool overview |  | Class A1 (Clinical use) | Class A2 (Troubleshooting) | Class A3 (Tube Replacement) | Class C1 (was in A) | Class C2 (was in C) | Class C3 (was in M) | Class M | protection level | Note | RT | Class A1 (Clinical use) | Class A2 (Basic Service) | Class A3 (Tube Replacement) | Class C1 (was in A) | Class C2 (was in C) | Class C3 (was in M) | Class M | protection level | Action Needed | Li Jinag Comment | Tool overview | 使ってみて
Error Logs (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
Storelog | ERROR_STORE_LOG | troubleshooting+tube replacement | archives the gesyslog and starts a blank copy |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  |  | archives the gesyslog and starts a blank copy | 超小型版、sprsnapみたいなlog収集ツール。
System Browser (All features ) | ERROR_SYSTEM_BROWSER | troubleshooting+tube replacement | Viewer for many files such as gesyslog, syslog, tube usage, software health page, gantry balance log….. |  |  |  |  | C1 | C1 | C1 | C1 | 1 |  | X |  |  |  | C1 | C1 | C1 | C1 | 1 |  |  | Viewer for many files such as gesyslog, syslog, tube usage, software health page, gantry balance log…..
Tube Usage | ERROR_TUBE_USAGE | troubleshooting+tube replacement | displays tube usage data |  |  |  |  | C1 | C1 | C1 | C1 | 1 |  | X |  |  |  | C1 | C1 | C1 | C1 | 1 |  |  | displays tube usage data
System Error Log | ERROR_SYSTEM_LOG | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 2
Software History Logs | ERROR_SOFTWARE_HISTORY_LOGS | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 3
StartGFV | ERROR_START_GFV | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 1
Diagnostics (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
kV & mA (X-ray) | DIAG_KV_MA | troubleshooting | diag utility to test kV and mA functionality |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 2 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 2 |  |  | diag utility to test kV and mA functionality | Revoでは、KV mA Scan Tool
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI | troubleshooting | JEDI specific utilities and diags |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 | Class C2 features (HHS Scan and HHS Report) are enabled by individual Token ID's. Class M feature (kV Diag with rotation) is enabled by Class M Service Key. | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  |  | JEDI specific utilities and diags |  | incl. HHS Scans and report,HHS in C2 level
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE | troubleshooting | allows manual exercising of collimator apertures |  |  |  |  | C1 | C1 | C1 | C1 | 3 |  | X |  |  |  | C1 | C1 | C1 | C1 | 4 |  |  | allows manual exercising of collimator apertures
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL | troubleshooting | utility to exercise the collimator apertures and filters for various gantry parameters and display positioning data |  |  |  |  | C1 | C1 | C1 | C1 | 2 |  | X |  |  |  | C1 | C1 | C1 | C1 | 2 |  |  | utility to exercise the collimator apertures and filters for various gantry parameters and display positioning data
Axial Functional | DIAG_AXIAL_FUNCTIONAL | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 3
DAS Tools | DIAG_DAS_TOOLS | troubleshooting | DAS/Detector integration tool that contains various tests |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4 |  |  | DAS/Detector integration tool that contains various tests
Scan Data Path | DIAG_SCAN_DATA_PATH | troubleshooting | Diag that tests the data chain from detector to host at defined stages |  |  |  |  | C1 | C1 | C1 | C1 | 2 |  | X |  |  |  | C1 | C1 | C1 | C1 | 2 |  |  | Diag that tests the data chain from detector to host at defined stages
Diagnatoc Data Collection (DDC) | DIAG_DATA_COLLECTION | troubleshooting | Scan data creation utility allowing configurable scan parameters bypassing certain lockouts for diagnostic use. |  |  |  |  | C1 | C1 | C1 | C1 | 3 |  | X |  |  |  | C1 | C1 | C1 | C1 | 4 |  |  | Scan data creation utility allowing configurable scan parameters bypassing certain lockouts for diagnostic use.
DIP Diagnostics | DIAG_DIP | troubleshooting | DIP diagnostic utility (subset of console diags) |  |  |  |  | C1 | C1 | C1 | C1 | 3 |  | X |  |  |  | C1 | C1 | C1 | C1 | 3 |  |  | DIP diagnostic utility (subset of console diags)
High Speed Disk Test | DIAG_HIGH_SPEED_DISK | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 3
Recon Data Path | DIAG_RECON_DATA_PATH | troubleshooting | Diag that test the recon data paths in the host |  |  |  |  | C1 | C1 | C1 | C1 | 2 |  | X |  |  |  | C1 | C1 | C1 | C1 | 2 |  |  | Diag that test the recon data paths in the host
scsistat OC | DIAG_SCSISTAT_OC | troubleshooting | Displays the scsi information from the host | [5 empty] | C2 | C2 | C2 | 1 |  | X |  |  |  |  | C2 | C2 | C2 | 1 | Remove? |  | Displays the scsi information from the host | unixのscsistatと同じ。
rhapSnap | DIAG_RHAP_SNAP | troubleshooting | sprsnap type utility |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 1 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 1 | Remove? |  | sprsnap type utility | 小型版、sprsnapみたいなlog収集ツール。
GSCB Keyboard Test | DIAG_GSCB_KEYBOARD | troubleshooting | Diag for the SCIM | [5 empty] | C2 | C2 | C2 | 2 |  | X |  |  |  |  | C2 | C2 | C2 | 2 |  | change to GSCB Keyboard Test after HDe6 software | Diag for the SCIM
RTS Log Viewer | DIAG_RTS_LOG_VIEWER | troubleshooting | GUI viewer for the RTS data collected by the system. | [9 empty] | Need to define the token name, class and protection level. | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA | Protection level required. |  | GUI viewer for the RTS data collected by the system.
RTS Viewer | DIAG_RTS_VIEWER | [12 empty] | X |  |  |  | C1 | C1 | C1 | C1 | 3
System Troubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | troubleshooting | Troubleshooting utility allowing viewing of gantry parameters for "limited access" installations where access to TGPU is restricted. |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | Troubleshooting utility allowing viewing of gantry parameters for "limited access" installations where access to TGPU is restricted. | Revoでは、別もので、System Cabinet Remote Managerがいる（SysCab rmmの実施と同じ） また、System Cabinet Diagnosticsもいる（SysCab diagの実施と同じ）。
Console Diagnostics | DIAG_CONSOLE | troubleshooting | suite of tools for testing host and recon hardware | [5 empty] | C2 | C2 | C2 | 3 |  | X |  |  |  |  | C2 | C2 | C2 | 3 |  |  | suite of tools for testing host and recon hardware
Smart kV Tool | DIAG_SMART_KV_TOOL | troubleshooting | Utility that allows scanning of a phantom to replace the HV Bleeder for kV testing. | [6 empty] | C3 | C3 | 3 |  | X | [5 empty] | C3 | C3 | 4 |  |  | Utility that allows scanning of a phantom to replace the HV Bleeder for kV testing.
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | [12 empty] | X | [6 empty] | M | 3
Beam Quality Scan | DIAG_BEAM_QUALITY_SCAN | [12 empty] | X | [6 empty] | M | 1
Image Quality (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
Image Analysis | IQ_IMAGE_ANALYSIS | troubleshooting | Utility that allows manual and automatic image analysis for IQ parameters. |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA | Same token but functionality is "Automated Image Analysis". | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | Utility that allows manual and automatic image analysis for IQ parameters.
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION | troubleshooting | Installs the Image orientation networking test patterns into the image database |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 3 | Changed to A1. HD is already A1. | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | same as Revolution. A2 for CJ, A1 for HD. Need to confirm why it's changed to A1. | Installs the Image orientation networking test patterns into the image database
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | clinical use+troubleshooting | Installs the SMPTE monitor test pattern into the image database |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | Installs the SMPTE monitor test pattern into the image database
Scan Analysis | IQ_SCAN_ANALYSIS | troubleshooting | Utility that allows viewing and analysis of scan and calibration data and viewing of header/footer information. |  |  |  |  | C1 | C1 | C1 | C1 | 2 |  | X |  |  |  |  | C1 | C1 | C1 | 2 | 2 |  |  | Utility that allows viewing and analysis of scan and calibration data and viewing of header/footer information.
Install TG18 Image | IQ_INSTALL_TG18_IMAGE | [12 empty] | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | DINのため。 |  | it is in toolchest
Calibration (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
Generator Filament Calibration | CAL_GENERATOR_FILAMENT | troubleshooting+tube replacement | Runs the generator filiment calibration tool |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  |  | Runs the generator filiment calibration tool
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION | troubleshooting | characterization tool for manually rotating the image to account for small errors between home position and image display |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4 |  |  | characterization tool for manually rotating the image to account for small errors between home position and image display
BOW Alignment | CAL_BOW_ALIGNMENT | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4
CBF and SAG Alignment | CAL_CBF_SAG_ALIGNMENT | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4
Collimator Cal | CAL_COLLIMATOR | troubleshooting+tube replacement | runs the collimator calibration tool |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 4 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4 |  |  | runs the collimator calibration tool
DAS Gain Cal | CAL_DAS_GAIN | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT | troubleshooting+tube replacement | runs the HOT ISO tool |  |  |  | A3 | A3 | A3 | A3 | A3 | 3 |  | X |  |  | A3 | A3 | A3 | A3 | A3 | 4 |  |  | runs the HOT ISO tool
ISO Alignment | CAL_ISO_ALIGNMENT | troubleshooting+tube replacement | runs the ISO tool |  |  |  | A3 | A3 | A3 | A3 | A3 | 3 |  | X |  |  | A3 | A3 | A3 | A3 | A3 | 4 |  |  | runs the ISO tool
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION | troubleshooting | Table characterization utility |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4 |  |  | Table characterization utility | Revoでは、Conigurationの中にいる。
Detailed Calibration | CAL_DETAILED | clinical use+troubleshooting+tube replacement | starts detailed calibration |  |  |  | A3 | A3 | A3 | A3 | A3 | 4 |  | X |  |  | A3 | A3 | A3 | A3 | A3 | 4 |  |  | starts detailed calibration
Z Slope Cal | CAL_Z_SLOPE | troubleshooting+tube replacement | GUI for aligning tube |  |  |  | A3 | A3 | A3 | A3 | A3 | 3 |  | X |  |  | A3 | A3 | A3 | A3 | A3 | 4 |  |  | GUI for aligning tube
POR Alignment | CAL_POR_ALIGNMENT | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4
Gantry Service Balance | CAL_GANTRY_SERVICE_BALANCE | troubleshooting+tube replacement | GUI for performing gantry balance |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 2 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 2 |  |  | GUI for performing gantry balance
RTP Alignment Tool | CAL_RTP_ALIGNMENT | [12 empty] | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
Configuration (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
Config Tracker | CONFIG_TRACKER | clinical use+troubleshooting | basic viewer listing (swhwinfo) |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 1 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 1 |  |  | basic viewer listing (swhwinfo)
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | clinical use+troubleshooting | Enhanced Xtream Injector config utility |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | Enhanced Xtream Injector config utility
Install Options | CONFIG_INSTALL_OPTIONS | troubleshooting | Option install utility |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | Option install utility
Product SW Revs | CONFIG_PRODUCT_SW_REVS | clinical use+troubleshooting | showprods type utility |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 1 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 1 |  |  | showprods type utility
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID | troubleshooting | Not used in VCT or HD |  |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 3 |  | This tool is deleted from HDe6 | Not used in VCT or HD
Install Camera | CONFIG_INSTALL_CAMERA | troubleshooting | DICOM interface install utility |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | DICOM interface install utility
Verify Options | CONFIG_VERIFY_OPTIONS | troubleshooting | displays options list that has been installed. |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | displays options list that has been installed.
Tube Install Certification | CONFIG_TUBE_INSTALL_CERTIFICATION | troubleshooting+tube replacement | GE proprietary tube install utility |  |  |  | A3 | A3 | A3 | A3 | A3 | 3 |  | X |  |  | A3 | A3 | A3 | A3 | A3 | 3 |  |  | GE proprietary tube install utility
Configure Enhanced PNF | CONFIG_ENH_PNF | clinical use+troubleshooting | utility to set up the firewall |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
Configure EA3 | CONFIG_EA3 |  | EA3 user registration. |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | utility to set up the firewall
Configure EAT | CONFIG_EAT |  | EAT user registration. |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | EA3 user registration.
Configure HIPAA | CONFIG_HIPAA |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | EAT user registration.
InSite Status | CONFIG_INSITE_STATUS | [12 empty] | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
Configure InSite | CONFIG_INSITE | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 1
Configure PRODIAGS | CONFIG_PRODIAGS | [12 empty] | X | [6 empty] | M | 1
Configure Software Download | CONFIG_SW_DOWNLOAD | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 1
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | clinical use+troubleshooting | AW network configuration utility |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
Configure AWS on VM | None | [12 empty] | X | X
AWS Client Install/Uninstall | CONFIG_AWVS_CLIENT_INSTALL | [12 empty] | X | X | A1 | A1 | A1 | A1 | A1 | A1 | NA | NA | A1 | NA
Configure AWS Applications | CONFIG_AWVS_APPLICATIONS | [12 empty] | X | X | A1 | A1 | A1 | A1 | A1 | A1 | NA | NA | A1 | NA
Launch AWS Service Tool | CONFIG_LAUNCH_AWVS_SERVICE | [12 empty] | X | X | A1 | A1 | A1 | A1 | A1 | A1 | NA | NA | A1 | NA
Configure AW Server Client | CONFIG_AW_SERVER_CLIENT |  | AW server configuration utility |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA | Token required. |  | AW network configuration utility
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | clinical use+troubleshooting | Standalone EKG monitor install utility |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | AW server configuration utility
Component Tracker | NA | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 1
Console | CONFIG_COMPONENT_TRACKER_CONSOLE | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 1
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 1
History | CONFIG_COMPONENT_TRACKER_HISTORY | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 1
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA | [12 empty] | X |  |  |  |  | C2 | C2 | C2 | 1
Config D4D | CONFIG_D4D | [13 empty] | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
Utilities (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | troubleshooting | Utility to download FW code to system components |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | Utility to download FW code to system components
HHS Scans | UTIL_HHS_SCANS | [17 empty] | C2 | C2 | C2 | 4
HHS Report | UTIL_HHS_REPORT | [17 empty] | C2 | C2 | C2 | 4
Auto CT# Adjust | CAL_WATER | [14 empty] | A2 | A2 | A2 | A2 | A2 | A2 | 4
Heat Soak and Seasoning | UTIL_HEAT_SOAK_SEASONING | [14 empty] | A2 | A2 | A2 | A2 | A2 | A2 | 3
Verify Security | UTIL_VERIFY_SECURITY | troubleshooting | check status of security key |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | check status of security key
Editor | UTIL_EDITOR | clinical use+troubleshooting | basic editor window "gedit" |  |  |  |  | C1 | C1 | C1 | C1 | 1 |  | X |  |  |  | C1 | C1 | C1 | C1 | 1 |  |  | basic editor window "gedit" | unixのxeditと同じ。 Revoでは、Calculatorが残っている。
Shell | UTIL_SHELL | clinical use+troubleshooting | opens unix shell "xterm" |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | opens unix shell "xterm"
Change Display IP | UTIL_CHANGE_DISPLAY_IP | clinical use+troubleshooting | video redirect for insite |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | video redirect for insite
TVA | UTIL_TVA | [12 empty] | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA | NA
Install Software Patch | UTIL_INSTALL_SW_PATCH | [12 empty] | X | [6 empty] | M | 1
Security Center | UTIL_SECURITY_CENTER | [12 empty] | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
System State - USB | UTIL_SYSTEM_STATE_USB |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | clinical use+troubleshooting | HIPAA setup utility, setting up user accounts |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | HIPAA setup utility, setting up user accounts
Reset Admin User Password | UTIL_RST_ADM_PASSWD | [19 empty] | M | 1
Reset Service User Password | UTIL_RST_SRV_PASSWD | [19 empty] | M | 1
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | clinical use+troubleshooting | apps shutdown |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  | apps shutdown
Audit Log Viewer | UTIL_AUDIT_LOG_VIEWER |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA
Replacement (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
New Tube Config | REP_NEW_TUBE_CONFIG | troubleshooting+tube replacement | Documenting new tube ID |  |  |  | A3 | A3 | A3 | A3 | A3 | 3 |  | X |  |  | A3 | A3 | A3 | A3 | A3 | 3 |  |  | Documenting new tube ID
Product Locator | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | NA |  |  |  | CSO対応で星野さんがcodeをチェックしてくれていた。 RevoでDICOM CSではDevice Serial Numberは、not sent, or blank。
Login | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1
History of Change | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1
Create Baseline | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1
Display Baseline | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1
Suite Information | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1
GIB Information | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1
Current Data | N/A |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1
PM (CSD menu) |  |  |  |  | * | * | * | * | * | * | *
Service Notepad | PM_SERVICE_NOTEPAD |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X |  |  |  |  | C2 | C2 | C2 | 3 |  |  | notepad utility for commenting into session log | Session Logに記録する機能だった。
Session Log | PM_SESSION_LOG |  |  |  | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 1 |  | X |  |  |  |  | C2 | C2 | C2 | 3
Scanner Utilities (Separated located at bottom of left monitor) |  |  |  |  | * | * | * | * | * | * | *
Detailed Calibration | CAL_DETAILED | [12 empty] | X |  |  | A3 | A3 | A3 | A3 | A3 | 4
Center Phantom | CAL_CENTER_PHANTOM | troubleshooting | notepad utility for commenting into session log | [5 empty] | C2 | C2 | C2 | 3 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 4
Z-Slope Calibration | CAL_Z_SLOPE | [12 empty] | X |  |  | A3 | A3 | A3 | A3 | A3 | 4
Manual CT # Adjust | UTIL_MANUAL_CT_ADJUST | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4
Auto CT # Adjust | CAL_WATER | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4
Collimator Calibration | CAL_COLLIMATOR | troubleshooting | Display system component status and break down to likely root cause when status is not "green". | [5 empty] | C2 | C2 | C2 | 3 |  | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4
Air Ratio Calibration | CAL_AIR_RATIO | troubleshooting | session log viewer | [5 empty] | C2 | C2 | C2 | 2 |  | X | A1 | A1 | A1 | A1 | A1 | A1 | A1 | 4
DAS Gain Cal | CAL_DAS_GAIN | [12 empty] | X |  | A2 | A2 | A2 | A2 | A2 | A2 | 4