## Optima CT520 Toolslist-TangX
Optima CT520 phase1.5 Service tools reclassification
System tools/utilities/diags | Tool License Token | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | protection level
Error Logs (CSD menu)
System Error Log | ERROR_SYSTEM_LOG |  |  |  |  | X |  |  | 2
Software History Log | ERROR_SOFTWARE_HISTORY_LOGS |  |  |  |  | X |  |  | 3
Storelog | ERROR_STORE_LOG |  | X | [5 empty] | 3
System Browser (All features ) | ERROR_SYSTEM_BROWSER |  |  |  | X |  |  |  | 1
Tube Usage | ERROR_TUBE_USAGE |  |  |  | X |  |  |  | 1
LogViewer | ERROR_LOG_VIEWER |  |  |  | X |  |  |  | 3
StartGFV | ERROR_START_GFV |  |  |  |  | X |  |  | 1
Diagnostics (CSD menu)
LAN Diagnostics | DIAG_LAN |  |  |  |  | X |  |  | 3
CPU Processes | DIAG_CPU_PROCESSES |  |  |  |  | X |  |  | 1
kV & mA (X-ray) | DIAG_KV_MA |  | X | [5 empty] | 2
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI |  | X |  |  | X |  | X | 4
Generator Tool – JEDI Filament Cal | CAL_GENERATOR_FILAMENT |  | X | [5 empty] | 4
Generator Tool – JEDI HHS Scans | UTIL_HHS_SCANS |  |  |  |  | X |  |  | 4
Generator Tool – JEDI HHS Report | UTIL_HHS_REPORT |  |  |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics Inverter Short Circuit | DIAG_GENERATOR_TOOL_JEDI |  | X |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics No Load KV | DIAG_GENERATOR_TOOL_JEDI |  | X |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics No Load KV(Rotation) | [7 empty] | X | 4
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE |  |  |  | X |  |  |  | 4
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL |  |  |  | X |  |  |  | 2
Axial Functional | DIAG_AXIAL_FUNCTIONAL |  |  |  |  | X |  |  | 3
DAS Tools | DIAG_DAS_TOOLS |  | X |  |  | X |  |  | 4
Scan Data Path | DIAG_SCAN_DATA_PATH |  |  |  | X |  |  |  | 2
Diagnostic Data Collection (DDC) | DIAG_DATA_COLLECTION |  |  |  | X |  |  |  | 4
DIP Diagnostics | DIAG_DIP |  |  |  | X |  |  |  | 1
High Speed Disk Test | DIAG_HIGH_SPEED_DISK |  |  |  |  | X |  |  | 1
Recon Data Path | DIAG_RECON_DATA_PATH |  |  |  | X |  |  |  | 2
scsistat OC | DIAG_SCSISTAT_OC |  |  |  |  | X |  |  | 1
rhapSnap | DIAG_RHAP_SNAP |  | X | [5 empty] | 1
GSCB Keyboard Test | DIAG_GSCB_KEYBOARD |  |  |  |  | X |  |  | 2
RTS Viewer (Basic) | DIAG_RTS_VIEWER |  |  |  | X |  |  |  | 3
RTS Viewer (Advanced) | DIAG_RTS_VIEWER_C |  |  |  |  | X |  |  | 4
System Toubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | X | [6 empty] | NA
Console Diagnostics | DIAG_CONSOLE |  |  |  |  | X |  |  | 1
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | [6 empty] | X | 3
Smart kV Tool | DIAG_SMART_KV_TOOL | [5 empty] | X |  | 4
Image Quality (CSD menu)
Image Analysis | IQ_IMAGE_ANALYSIS | X | [6 empty] | NA
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | X | [6 empty] | NA
Scan Analysis | IQ_SCAN_ANALYSIS |  |  |  | X |  |  |  | 2
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION |  | X | [5 empty] | 3
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES | X | [6 empty] | NA
Calibration (CSD menu)
BOW Alignment |  |  | x | [5 empty] | 4
Generator Filament Calibration | CAL_GENERATOR_FILAMENT |  | X | [5 empty] | 4
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION |  | X | [5 empty] | 4
CBF and SAG Alignment |  |  | x | [5 empty] | 4
Collimator Cal | CAL_COLLIMATOR |  | X | [5 empty] | 4
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT |  |  | X |  |  |  |  | 4
ISO Alignment | CAL_ISO_ALIGNMENT |  |  | X |  |  |  |  | 4
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION |  | X | [5 empty] | 4
Default Patient Positioning | CAL_DEFAULT_PATIENT_POSITIONING | X | [6 empty] | NA
POR Alignment | CAL_POR_ALIGNMENT |  | x | [5 empty] | 3
Air cal(if no cals loaded) | NA | x | [6 empty] | 4
Detailed Calibration | CAL_DETAILED |  | X | [5 empty] | 4
Configuration (CSD menu)
Config Tracker | CONFIG_TRACKER |  | X | [5 empty] | 1
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | X | [6 empty] | NA
Install Options | CONFIG_INSTALL_OPTIONS | X | [6 empty] | NA
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID |  | X | [5 empty] | 3
Install Camera | CONFIG_INSTALL_CAMERA | X | [6 empty] | NA
Verify Options | CONFIG_VERIFY_OPTIONS | X | [6 empty] | NA
Tube Install Certification with Non-GE tube | CONFIG_TUBE_INSTALL_CERTIFICATION |  |  | X |  |  |  |  | 3
Tube Install Certification with GE Tube | CONFIG_TUBE_INSTALL_CERTIFICATION | [6 empty] | X | 3
Configure PNF | CONFIG_PNF | X | [6 empty] | NA
Configure EA3 | CONFIG_EA3 | X | [6 empty] | NA
Configure EAT | CONFIG_EAT | X | [6 empty] | NA
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | X | [6 empty] | NA
Configure AW Server Client | NA | X | [6 empty] | NA
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | X | [6 empty] | NA
Component Tracker | NA
Axial | CONFIG_COMPONENT_TRACKER_AXIAL |  |  |  |  | X |  |  | 1
Console | CONFIG_COMPONENT_TRACKER_CONSOLE |  |  |  |  | X |  |  | 1
Detector | CONFIG_COMPONENT_TRACKER_DETECTOR |  |  |  |  | X |  |  | 1
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE |  |  |  |  | X |  |  | 1
History | CONFIG_COMPONENT_TRACKER_HISTORY |  |  |  |  | X |  |  | 1
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA |  |  |  |  | X |  |  | 1
Patient Information on Xtream Display | N/A | X | [6 empty] | NA
InSite Status | CONFIG_INSITE_STATUS | X | [6 empty] | NA
Configure InSite | CONFIG_INSITE |  |  |  |  | X |  |  | 3
Configure PRODIAGS | CONFIG_PRODIAGS | [6 empty] | X | 2
Configure Software Download | CONFIG_SW_DOWNLOAD |  | X | [5 empty] | 2
Configure AWS on VM | None
AWS Client Install/Uninstall | CONFIG_AWVS_CLIENT_INSTALL | X | [6 empty] | NA
Configure AWS Applications | CONFIG_AWVS_APPLICATIONS | X | [6 empty] | NA
Launch AWS Service Tool | CONFIG_LAUNCH_AWVS_SERVICE | X | [6 empty] | NA
Gantry Table User Preference | CONFIG_GANTRY_TABLE_PREFS | X | [6 empty] | NA
Configure HIPAA | CONFIG_HIPAA | X | [6 empty] | NA
Utilities (CSD menu)
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | X | [6 empty] | NA
HHS Scans(Class A manual exist method) | UTIL_HHS_SCANS |  |  |  |  | X |  |  | 3
HHS Report | UTIL_HHS_REPORT |  |  |  |  | X |  |  | 3
Manual CT # Adjust | UTIL_MANUAL_CT_ADJUST |  | X | [5 empty] | 4
Auto CT # Adjust | CAL_WATER |  | X | [5 empty] | 4
Heat Soak and Seasoning (A&C are same) | UTIL_HEAT_SOAK_SEASONING |  | X | [5 empty] | 3
CT Mobile | UTIL_CT_MOBILE | X | [6 empty] | NA
Verify Security | UTIL_VERIFY_SECURITY | X | [6 empty] | NA
Editor | UTIL_EDITOR |  |  |  | X |  |  |  | 1
Calculator | UTIL_CALCULATOR |  |  |  | X |  |  |  | 1
Shell | UTIL_SHELL | X | [6 empty] | NA
Change Display IP | UTIL_CHANGE_DISPLAY_IP | X | [6 empty] | NA
System State（removed in TangX） | UTIL_SYSTEM_STATE | X
System State - USB | UTIL_SYSTEM_STATE_USB | X | [6 empty] | NA
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | X | [6 empty] | NA
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | X | [6 empty] | NA
TVA | UTIL_TVA | X | [6 empty] | NA
Install Software Patch | UTIL_INSTALL_SW_PATCH | [6 empty] | X | 1
Security Center | UTIL_SECURITY_CENTER | X | [6 empty] | NA
Audit Log Viewer | UTIL_AUDIT_LOG_VIEWER | X | [6 empty] | NA
Replacement (CSD menu)
New Tube Config | REP_NEW_TUBE_CONFIG |  |  | X |  |  |  |  | 3
FRDM Wizard | REP_FRDM_WIZARD |  | X | [5 empty] | 1
Product Locator | N/A | X | [6 empty] | NA
PM (CSD menu)
Service Notepad | PM_SERVICE_NOTEPAD |  |  |  |  | X |  |  | 3
Session Log | PM_SESSION_LOG |  |  |  |  | X |  |  | 2
System Health Page | PM_SYSTEM_HEALTH_PAGE |  |  |  |  | X |  |  | 3
Scan Utilities
Scan Utilties - Detailed Carlibration | CAL_DETAILED |  | x | [5 empty] | 4
Scan Utilties - Center Phantom | CAL_CENTER_PHANTOM | X | [6 empty] | 4
Scan Utilties - Auto CT Number Adjust | CAL_WATER |  | x | [5 empty] | 4
Scan Utilties - Collimator Cal | CAL_COLLIMATOR |  | x | [5 empty] | 4
Scan Utilties - Air Calibration | CAL_AIR | X | [6 empty] | 4
Scan Utilties - Spectral Cal | CAL_SPECTRAL |  |  | X |  |  |  |  | 4