## Cj Phase3&KL64ToolsList (hidden)
Cj Ph3 Service tools reclassification Final Oct. 11th., 2013 | [19 empty] | Kunlun 16
view Service Document |  |  |  | Current IB Classification |  |  |  | New IB Classification - Cj | [11 empty] | view Service Document |  |  |  | Current IB Classification |  |  |  | New IB Classification - Cj
System tools/utilities/diags (As they exist on VCT/HD today, very similar to older CT IB systems as well) | Tool License Token |  | Tool overview |  | Class A | Class C | Class M |  | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | protection level | Impact of reclassification to Indepedant Service Provider (Class A moved to Class C) | Comments | Blank | System tools/utilities/diags (As they exist on KL16 today, very similar to older CT IB systems as well) | Tool License Token | CJ tools name | Tool overview |  | Class A | Class C | Class M |  | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | Protection Level | Comments
Error Logs (CSD menu) | [19 empty] | Error Logs (CSD menu)
System Error Log | ERROR_SYSTEM_LOG | troubleshooting+tube replacement | pulls up gesyslog in unix shell "view" type output |  |  | X | [6 empty] | X |  |  | 2 |  | No protection to log files anyway. And it's using a viewlog command which may be widely used. |  | System Error Log | ERROR_SYSTEM_LOG | System Error Log |  |  |  | X | [6 empty] | X |  |  | 2 | No protection to log files anyway. And it's using a viewlog command which may be widely used.
Software History Log | ERROR_SOFTWARE_HISTORY_LOGS | troubleshooting+tube replacement | GUI allowing user to view all the history logs for tools and diags |  |  | X | [6 empty] | X |  |  | 3 |  |  |  | Software History Logs | ERROR_SOFTWARE_HISTORY_LOGS | Software History Log |  |  |  | X | [6 empty] | X |  |  | 3
Storelog | ERROR_STORE_LOG | troubleshooting+tube replacement | archives the gesyslog and starts a blank copy |  | X |  |  |  |  | X | [5 empty] | 3 |  |  |  | Store Log | ERROR_STORE_LOG | Storelog |  |  | X |  |  |  |  | X | [5 empty] | 3
System Browser (All features ) | ERROR_SYSTEM_BROWSER | troubleshooting+tube replacement | Viewer for many files such as gesyslog, syslog, tube usage, software health page, gantry balance log….. |  | X | [6 empty] | X |  |  |  | 1 | Can use unix editor to view log files | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param. |  | System Browser(All feature) | ERROR_SYSTEM_BROWSER | System Browser (All features ) |  |  | X | [6 empty] | X |  |  |  | 1 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
Tube Usage | ERROR_TUBE_USAGE | troubleshooting+tube replacement | displays tube usage data |  | X | [6 empty] | X |  |  |  | 1 | Tool used specifically for GE warranty tracking. Not required if not GE serviced. | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param. |  | Tube Usage | ERROR_TUBE_USAGE | Tube Usage |  |  | X | [6 empty] | X |  |  |  | 1 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
LogViewer | ERROR_LOG_VIEWER | troubleshooting+tube replacement | terra log viewer, shows arslog, dbrlog, browser log, net log…… |  | X | [6 empty] | X |  |  |  | 3 | Can use unix editor to view log files |  |  | LogViewer | ERROR_LOG_VIEWER | LogViewer |  |  | X | [6 empty] | X |  |  |  | 3
StartGFV | ERROR_START_GFV | troubleshooting+tube replacement | GE File Viewer for gesyslog with pattern detection. |  |  | X | [6 empty] | X |  |  | 1 |  | Technical difficulty: it's only a tar package installed to console. |  | StartGFV | ERROR_START_GFV | StartGFV |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty: it's only a tar package installed to console.
Diagnostics (CSD menu) | [19 empty] | Diagnostics (CSD menu)
LAN Diagnostics | DIAG_LAN | troubleshooting | ping tool to OC, TGP, ORP, DARC |  |  | X | [6 empty] | X |  |  | 3 |  |  |  | LAN Diagnositics | DIAG_LAN | LAN Diagnostics |  |  |  | X | [6 empty] | X |  |  | 3
CPU Processes | DIAG_CPU_PROCESSES | troubleshooting | windows task manager concept |  |  | X | [6 empty] | X |  |  | 1 |  | This is a Linux built-in command. No need to protect it. |  | CPU Processes | DIAG_CPU_PROCESSES | CPU Processes |  |  |  | X | [6 empty] | X |  |  | 1 | This is a Linux built-in command. No need to protect it.
kV & mA (X-ray) | DIAG_KV_MA | troubleshooting | diag utility to test kV and mA functionality |  | X |  |  |  |  | X | [5 empty] | 2 | Currently contains functions to collected data for regulatory filings not just diagnostics so need to remain class A. | Technical difficulty to protect the java package because it's shared by many programs. |  | kV & mA (X-ray) | DIAG_KV_MA | kV & mA (X-ray) |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI | troubleshooting | JEDI specific utilities and diags |  | X | X | X |  |  | X |  |  | X |  | X | 3 | Currently contains functions to collected data for regulatory filings not just diagnostics so need to remain class A. KV Diagnostics (No Load KV and Inverter short circuit) can be Class C. | Program file is only one. But the functionality should be changed by Class C/M. |  | Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI | Generator Tool - JEDI |  |  | X | X | X |  |  | X |  |  | X |  | X | 3 | Need to include KVMA in license
Not for Cj M40 and KL16 | DIAG_FOCAL_SPOT | [18 empty] | Generator Tool JEDI - Focal Spot | DIAG_FOCAL_SPOT | [14 empty] | 4 | Need classification.
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE | troubleshooting | allows manual exercising of collimator apertures |  | X | [6 empty] | X |  |  |  | 3 | Test helps isolate internal to collimator. Replacement of collimator also resolves issue. |  |  | Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE | Collimator Aperture Test |  |  | X | [6 empty] | X |  |  |  | 3
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL | troubleshooting | utility to exercise the collimator apertures and filters for various gantry parameters and display positioning data |  | X | [6 empty] | X |  |  |  | 2 | Test helps isolate internal to collimator. Replacement of collimator also resolves issue. | Technical difficulty to protect the java package because it's shared by many programs. |  | Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL | Collimator Functional |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Axial Functional | DIAG_AXIAL_FUNCTIONAL | troubleshooting | Utility that exercises the axial control function and displays actual vs expected data |  |  | X | [6 empty] | X |  |  | 3 | Test helps isolate axial drive issue giving more data to help troubleshoot. Does not point to exact FRU, just supplies more data. Replacement of multiple FRU's would resolve issue. |  |  | Axial Functional | DIAG_AXIAL_FUNCTIONAL | Axial Functional |  |  |  | X | [6 empty] | X |  |  | 3
DAS Tools | DIAG_DAS_TOOLS | troubleshooting | DAS/Detector integration tool that contains various tests |  | X | X |  |  |  | X |  |  | X |  |  | 3 | (QIF Visualization tool is Class C) |  |  | DAS Tools | DIAG_DAS_TOOLS | DAS Tools |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
DAS Tools - QIF Vis | DIAG_DAS_TOOLS_QIF_VIS | [11 empty] | X | [6 empty] | DAS Tools - QIF Vis | DIAG_DAS_TOOLS_QIF_VIS |  |  |  |  | X | [9 empty] | 4 | Need classification.
Scan Data Path | DIAG_SCAN_DATA_PATH | troubleshooting | Diag that tests the data chain from detector to host at defined stages |  | X | [6 empty] | X |  |  |  | 2 | Test helps isolate intermittant issue in scan data communication path. Replacement of multiple FRU's would resolve issue. | Technical difficulty to protect the java package because it's shared by many programs. |  | Scan Data Path | DIAG_SCAN_DATA_PATH | Scan Data Path |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Diagnatoc Data Collection (DDC) | DIAG_DATA_COLLECTION | troubleshooting | Scan data creation utility allowing configurable scan parameters bypassing certain lockouts for diagnostic use. |  | X | [6 empty] | X |  |  |  | 3 | Tool that allows easy configuration of scan parameters for troubleshooting outside the patient user interface through a simple GUI that also allows scans without xray. Can prescribe scans manually through the patient interface with xray. |  |  | Diagnostic Data Collection (DDC) | DIAG_DATA_COLLECTION | Diagnatoc Data Collection (DDC) |  |  | X | [6 empty] | X |  |  |  | 3 | DDC tool has problem with L3 protection, so changed to Level 4.
VDIP Diagnostics | DIAG_DIP | troubleshooting | DIP diagnostic utility (subset of console diags) |  | X | [6 empty] | X |  |  |  | 3 | Test helps isolate issue in scan data communication path within console. Replacement of multiple FRU's would resolve issue. | This is new tool in kl64 |  | DIP Diagnostics | DIAG_DIP | VDIP Diagnostics |  |  | X | [6 empty] | X |  |  |  | 3
High Speed Disk Test | DIAG_HIGH_SPEED_DISK | troubleshooting | Test utility for JBOD and HSDA disk drives |  |  | X | [6 empty] | X |  |  | 3 |  |  |  | High Speed Disk Test | DIAG_HIGH_SPEED_DISK | High Speed Disk Test |  |  |  | X | [6 empty] | X |  |  | 3
Recon Data Path | DIAG_RECON_DATA_PATH | troubleshooting | Diag that test the recon data paths in the host |  | X | [6 empty] | X |  |  |  | 2 | Test helps isolate issue in recon data communication path within console. Replacement of multiple FRU's would resolve issue. | Technical difficulty to protect the java package because it's shared by many programs. |  | Recon Data Path | DIAG_RECON_DATA_PATH | Recon Data Path | [9 empty] | x |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
VRAC Diagnostics | DIAG_RAC | troubleshooting | IG RAC card diagnostic utility (subset of console diags) |  |  | X | [6 empty] | X |  |  |  | only commands |  |  | RAC Diagnostics | DIAG_RAC | VRAC Diagnostics | [13 empty] | 3
scsistat OC | DIAG_SCSISTAT_OC | troubleshooting | Displays the scsi information from the host |  |  | X | [6 empty] | X |  |  | 1 |  |  |  | scsistat OC | DIAG_SCSISTAT_OC | scsistat OC |  |  |  | X | [6 empty] | X |  |  | 1 | This command is used by software startup, cannot protect it.
rhapSnap | DIAG_RHAP_SNAP | troubleshooting | sprsnap type utility |  |  | X |  |  |  | X | [5 empty] | 1 |  |  |  | rhapSnap | DIAG_RHAP_SNAP | rhapSnap |  |  |  | X | [6 empty] | X |  |  | 3
SCIM Keyboard Test | DIAG_GSCB_KEYBOARD | troubleshooting | Diag for the SCIM |  |  | X | [6 empty] | X |  |  | 2 |  | This is new tool in kl64(Technical difficulty to protect the java package.) |  | GSCB Keyboard Test | DIAG_GSCB_KEYBOARD | SCIM Keyboard Test |  |  |  | X | [6 empty] | X |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
RTS Viewer (class A) | DIAG_RTS_VIEWER | troubleshooting | GUI viewer for the RTS data collected by the system. |  | X | [6 empty] | X |  |  |  | 3 | GUI viewer of parametric data from system components. Can manually view data logs using unix commands to view a text log file. | class A is restrictive to just single exam data |  | RTS viewer | DIAG_RTS_VIEWER | RTS Viewer A |  |  | X | [6 empty] | X |  |  |  | 3
RTS Viewer (class C) | DIAG_RTS_VIEWER_C | [11 empty] | X |  |  | 4 |  | Class C mode allows cumulative statistics across many exams |  | RTS viewer C | DIAG_RTS_VIEWER_C | RTS Viewer C |  |  |  | X | [6 empty] | X |  |  | 4
System Toubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | troubleshooting | Troubleshooting utility allowing viewing of gantry parameters for "limited access" installations where access to TGPU is restricted. |  | X |  |  |  | X | [6 empty] | NA | easier to troubleshoot but can use mirror | Small room BCAT# B7700SB requires the tool at calibration. Then this tool should be class A. |  | System Troubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | System Toubleshooting Tool |  |  | X |  |  |  | X
Console Diagnostics | DIAG_CONSOLE | troubleshooting | suite of tools for testing host and recon hardware |  |  | X | [6 empty] | X |  |  | 3 |  |  |  | Console Diagnostics | DIAG_CONSOLE | Console Diagnostics |  |  |  | X | [6 empty] | X |  |  | 3 | Need to include DIP Diag in license
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | troubleshooting | Utility that displays system status and pattern recognition for diagnosis of system status/error reporting. |  |  |  | X | [7 empty] | X | 3 | Not worth to move to C as not supported and risk of compalints if moved to C |  |  | System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | System Health Monitor |  |  |  |  | X | [7 empty] | X | 3 | Whether we can wrap /vobs/srv/Tools/shm/UI/shmUI/dist/shmui.jar too?
Smart kV Tool | DIAG_SMART_KV_TOOL | troubleshooting | Utility that allows scanning of a phantom to replace the HV Bleeder for kV testing. |  |  |  | X | [6 empty] | X |  | 3 | Saving time : depending onnthe skills ; 1h15min to do bleader and 5min to do smart KV- non inviasive and no need to remove cover, not messy otherwhise needs to put oils…service procedure in class M high voltage tank feed-back resister verification & "restricited for smart kv" |  |  | Smart kV Tool | DIAG_SMART_KV_TOOL | Smart kV Tool |  |  |  |  | X | [6 empty] | X |  | 3
LB Tube Rotor Control (Class A) | DIAG_LB_TUBE_ROTOR_CONTROL |  |  |  | X |  |  |  |  | X | [5 empty] | 2 |  | This is new tool in kl64(Technical difficulty to protect the java package.)
LB Tube Rotor Control (Class C) | DIAG_LB_TUBE_ROTOR_CONTROL_C |  |  |  |  | X | [6 empty] | X |  |  | 2
LB Tube Rotor Phase Check | DIAG_LB_TUBE_ROTOR_PHASE_CHECK | [5 empty] | X | [7 empty] | X | 2 |  | This is new tool in kl64(Technical difficulty to protect the java package.)
Image Quality (CSD menu) | [19 empty] | Image Quality (CSD menu)
Image Analysis | IQ_IMAGE_ANALYSIS | troubleshooting | Utility that allows manual and automatic image analysis for IQ parameters. |  | X |  |  |  | X | [6 empty] | NA |  | Only A1 is considered in Japan, we would like to have it in A1. Suggested it in Tech Ref manual and User manual |  | Image Analysis | IQ_IMAGE_ANALYSIS | Image Analysis |  |  | X |  |  |  | X
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | clinical use+troubleshooting | Installs the SMPTE monitor test pattern into the image database |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | Install SMPTE Image |  |  | X |  |  |  | X
Scan Analysis | IQ_SCAN_ANALYSIS | troubleshooting | Utility that allows viewing and analysis of scan and calibration data and viewing of header/footer information. |  | X | [6 empty] | X |  |  |  | 2 | Utility that allows plotting of scan data to help figure out IQ issue root cause for subtle artifacts. Can replace multiple parts to resolve issue. | Technical difficulty to protect the java package. |  | Scan Analysis | IQ_SCAN_ANALYSIS | Scan Analysis |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package.
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION | troubleshooting | Installs the Image orientation networking test patterns into the image database |  | X |  |  |  |  | X | [5 empty] | 3 |  |  |  | Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION | Install Image Orientation Imgs |  |  | X |  |  |  |  | X | [5 empty] | 3 | No protection to the images.
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES |  |  |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES |  |  |  | X |  |  |  |  | X | [5 empty] | 3 | No protection to the images.
Calibration (CSD menu) | [19 empty] | Calibration (CSD menu)
Generator Filament Calibration | CAL_GENERATOR_FILAMENT | troubleshooting+tube replacement | Runs the generator filiment calibration tool |  | X |  |  |  |  | X | [5 empty] | 3 |  |  |  | Generator Filament Calibration | CAL_GENERATOR_FILAMENT | Generator Filament Calibration |  |  | X |  |  |  |  | X | [5 empty] | 3
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION | troubleshooting | characterization tool for manually rotating the image to account for small errors between home position and image display |  | X |  |  |  |  | X | [5 empty] | 3 |  |  |  | Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION | Gantry Rotation Characterization |  |  | X |  |  |  |  | X | [5 empty] | 3
Collimator Cal | CAL_COLLIMATOR | troubleshooting+tube replacement | runs the collimator calibration tool |  | X |  |  |  |  | X | [5 empty] | 4 |  |  |  | Collimator Cal | CAL_COLLIMATOR | Collimator Cal |  |  | X |  |  |  |  | X | [5 empty] | 4
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT | troubleshooting+tube replacement | runs the HOT ISO tool |  | X | [5 empty] | X |  |  |  |  | 3 |  |  |  | Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT | Hot ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 3
ISO Alignment | CAL_ISO_ALIGNMENT | troubleshooting+tube replacement | runs the ISO tool |  | X | [5 empty] | X |  |  |  |  | 3 |  |  |  | ISO Alignment | CAL_ISO_ALIGNMENT | ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 3
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION | troubleshooting | Table characterization utility |  | X |  |  |  |  | X | [5 empty] | 3 |  |  |  | Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION | Mechanical Characterization |  |  | X |  |  |  |  | X | [5 empty] | 3
Default Patient Positioning | CAL_DEFAULT_PATIENT_POSITIONING |  | DPP setting tool. |  | X |  |  |  | X | [6 empty] | NA | This is used for customer. I think this should be class A. |  |  | Default Patient Positioning | CAL_DEFAULT_PATIENT_POSITIONING | Default Patient Positioning |  |  | X |  |  |  | X
Detailed Calibration | CAL_DETAILED | clinical use+troubleshooting+tube replacement | starts detailed calibration |  | X | [5 empty] | X |  |  |  |  | 4 |  |  |  | Detailed Calibration | CAL_DETAILED | Detailed Calibration |  |  | X | [5 empty] | X |  |  |  |  | 4
BOW Alignment | CAL_BOW_ALIGNMENT | [16 empty] | CJ does not have this tool |  | BOW Alignment | CAL_BOW_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
POR Alignment | CAL_POR_ALIGNMENT | [16 empty] | CJ does not have this tool |  | POR Alignment | CAL_POR_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
CBF and SAG Alignment | CAL_CBF_SAG_ALIGNMENT | [16 empty] | CJ does not have this tool |  | CBF and SAG Alignment | CAL_CBF_SAG_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
Z Align | CAL_BOW_ALIGNMENT | troubleshooting+tube replacement | GUI for aligning tube |  | X | [5 empty] | X |  |  |  |  | 3 |  | This is new tool in kl64 |  |  | CAL_Z_ALIGN | Z Align |  |  | X |  |  |  |  | X | [5 empty] | TBD | Tool not used by KL16, TBD for CJ.
Smart Tube Alignment | CAL_SMART_TUBE_ALIGNMENT | troubleshooting+tube replacement | GUI for aligning tube |  |  |  | X | [6 empty] | X |  | 3 | it is a productivty tool. Value = combines 2 tools in one and gives the result at one time of the combination ( combines Z alignmenet and iso alignement) = time saving - 30min to 1h1/2 of time saving depending on skills. ( 2 direcitons on tubes and 2 bold - 1 tube does 1 direction and adjsut bold) - does separatly = 1h 15min vs 45min | This is new tool in kl64 |  |  | CAL_SMART_TUBE_ALIGNMENT | Smart Tube Alignment |  |  |  |  | X | [6 empty] | X |  | TBD | Tool not used by KL16, TBD for CJ.
Gantry Service Balance | CAL_GANTRY_SERVICE_BALANCE | troubleshooting+tube replacement | GUI for performing gantry balance |  | X |  |  |  |  | X | [5 empty] | 2 |  | This is new tool in kl64 |  |  | CAL_GANTRY_SERVICE_BALANCE | Gantry Service Balance |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package.
Scan Utilties - Center Phantom | CAL_CENTER_PHANTOM | troubleshooting+tube replacement | starts transfer function calibration |  | X |  |  |  | X | [6 empty] | 4 |  | For Kunlun 16, it is in Scan Utilties - Center Phantom |  | Scan Utilties - Center Phantom | CAL_CENTER_PHANTOM |  |  |  | x |  |  |  | x | [6 empty] | 4 | This is a button on Scan utilities.
Scan Utilties - Z Slope Cal | CAL_Z_SLOPE | [16 empty] | CJ does not have this tool |  | Scan Utilties - Z Slope Cal | CAL_Z_SLOPE |  |  |  | x | [5 empty] | X |  |  |  |  | 4 | This is a button on Scan utilities.
Scan Utilties - DAS Gain Cal | CAL_DAS_GAIN | [16 empty] | CJ does not have this tool |  | Scan Utilties - DAS Gain Cal | CAL_DAS_GAIN |  |  |  | x | [5 empty] | x |  |  |  |  | 4 | This is a button on Scan utilities.
Scan Utilties - Air Cal | CAL_AIR | [7 empty] | X | [6 empty] | 4 |  | For Kunlun 16, it is in Scan Utilties - Air Cal |  | Scan Utilties - Air Cal | CAL_AIR |  |  |  | x |  |  |  | x | [6 empty] | 4 | This is a button on Scan utilities.
Scan Utilties - Spectral Cal | CAL_SPECTRAL | [9 empty] | X |  |  |  |  | 4 |  | For Kunlun 16, it is in Scan Utilties - Spectral Cal |  | Scan Utilties - Spectral Cal | CAL_SPECTRAL |  |  |  | x | [5 empty] | x |  |  |  |  | 4 | This is a button on Scan utilities.
Configuration (CSD menu) | [19 empty] | Configuration (CSD menu) |  | Transfer Function Calibration
Config Tracker | CONFIG_TRACKER | clinical use+troubleshooting | basic viewer listing (swhwinfo) |  | X |  |  |  |  | X | [5 empty] | 1 |  | This tool is just a display of config info (which is generated after software LFC). Currently we cannot protect it. |  | Config Tracker | CONFIG_TRACKER | Config Tracker |  |  | X |  |  |  |  | X | [5 empty] | 1 | This tool is just a display of config info (which is generated after software LFC). Currently we cannot protect it.
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | clinical use+troubleshooting | Enhanced Xtream Injector config utility |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | Injector: Pressure Limit Unit |  |  | X |  |  |  | X
Install Options | CONFIG_INSTALL_OPTIONS | troubleshooting | Option install utility |  | X |  |  |  | X | [6 empty] | NA |  | should be A1, same as kl16 |  | Install Options | CONFIG_INSTALL_OPTIONS | Install Options |  |  | X |  |  |  | X
Product SW Revs | CONFIG_PRODUCT_SW_REVS | clinical use+troubleshooting | showprods type utility |  | X |  |  |  |  | X | [5 empty] | 1 |  | This command is used by software startup, cannot protect it. |  | Product SW Revs | CONFIG_PRODUCT_SW_REVS | Product SW Revs |  |  | X |  |  |  |  | X | [5 empty] | 1 | This command is used by software startup, cannot protect it.
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID | troubleshooting | Not used in VCT or HD |  | X |  |  |  |  | X | [5 empty] | 3 |  |  |  | Enter Detector ID | CONFIG_ENTER_DETECTOR_ID | Enter Detector ID |  |  | X |  |  |  |  | X | [5 empty] | 3
Install Camera | CONFIG_INSTALL_CAMERA | troubleshooting | DICOM interface install utility |  | X |  |  |  | X | [6 empty] | NA |  | should be A1, same as kl16 |  | Install Camera | CONFIG_INSTALL_CAMERA | Install Camera |  |  | X |  |  |  | X
Verify Options | CONFIG_VERIFY_OPTIONS | troubleshooting | displays options list that has been installed. |  | X |  |  |  | X | [6 empty] | NA |  | should be A1, same as kl16 |  | Verify Options | CONFIG_VERIFY_OPTIONS | Verify Options |  |  | X |  |  |  | X
Tube Install Certification | CONFIG_TUBE_INSTALL_CERTIFICATION | troubleshooting+tube replacement | GE proprietary tube install utility |  | X |  | X |  |  |  | X |  |  |  | X | 3 | contains class A functions as well as class M that are enabled by service key. |  |  | Tube Install Certification | CONFIG_TUBE_INSTALL_CERTIFICATION | Tube Install Certification |  |  | X |  | X |  |  |  | X |  |  |  | X | 3
Configure PNF | CONFIG_PNF | clinical use+troubleshooting | utility to set up the firewall |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Configure PNF | CONFIG_PNF | Configure PNF |  |  | X |  |  |  | X
Configure EA3 | CONFIG_EA3 |  | EA3 user registration. |  | X |  |  |  | X | [6 empty] | NA | I added and I think this is common with VCT. |  |  | Configure EA3 | CONFIG_EA3 | Configure EA3 |  |  | X |  |  |  | X
Configure EAT | CONFIG_EAT |  | EAT user registration. |  | X |  |  |  | X | [6 empty] | NA | I added and I think this is common with VCT. |  |  | Configure EAT | CONFIG_EAT | Configure EAT |  |  | X |  |  |  | X
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | clinical use+troubleshooting | AW network configuration utility |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | Configure AW DirectConnect |  |  | X |  |  |  | X
Configure AW Server Client |  |  | AW server configuration utility |  | X |  |  |  | X | [6 empty] | NA | I added and I think this is common with VCT. |  |  |  |  | Configure AW Server Client | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | clinical use+troubleshooting | Standalone EKG monitor install utility |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | Install Gating Monitor |  |  | X |  |  |  | X
Component Tracker |  | troubleshooting | [14 empty] | Provide history background, can come back to check what was done on CT- config history |  |  | Config Tracker |  | Config Tracker
Axial | CONFIG_COMPONENT_TRACKER_AXIAL |  |  |  |  | X | [6 empty] | X |  |  | 1 |  | This is new tool in kl64(Technical difficulty to protect such payload in the form of servlet.)
Console | CONFIG_COMPONENT_TRACKER_CONSOLE | troubleshooting | console hart data viewer |  |  | X | [6 empty] | X |  |  | 1 |  | Technical difficulty to protect such payload in the form of servlet. |  | Console | CONFIG_COMPONENT_TRACKER_CONSOLE | Console |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
Detector | CONFIG_COMPONENT_TRACKER_DETECTOR | troubleshooting | DAS/Detector hart data viewer |  |  | X | [6 empty] | X |  |  | 1 |  | This is new tool in kl64(Technical difficulty to protect such payload in the form of servlet.) |  |  | CONFIG_COMPONENT_TRACKER_DETECTOR | Detector | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE | troubleshooting | Software hart data viewer |  |  | X | [6 empty] | X |  |  | 1 |  | Technical difficulty to protect such payload in the form of servlet. |  | Software | CONFIG_COMPONENT_TRACKER_SOFTWARE | Software |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
History | CONFIG_COMPONENT_TRACKER_HISTORY | troubleshooting | History of changes to HART tracking files |  |  | X | [6 empty] | X |  |  | 1 |  | Technical difficulty to protect such payload in the form of servlet. |  | History | CONFIG_COMPONENT_TRACKER_HISTORY | History |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA | troubleshooting | Trigger a console hart data read |  |  | X | [6 empty] | X |  |  | 1 |  | Technical difficulty to protect such payload in the form of servlet. |  | Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA | Refresh Console Data |  |  |  | X | [6 empty] | X |  |  | 3
Patient Information on Xtream Display | N/A |  | Xtream Display management. |  | X |  |  |  | X | [6 empty] | NA | Cj unique. | This is new tool in kl64(Technical difficulty to protect such payload in the form of servlet.) |  |  |  | Patient Information on Xtream Display | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
Utilities (CSD menu) | [19 empty] | Utilities (CSD menu)
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | troubleshooting | Utility to download FW code to system components |  | X |  |  |  | X | [6 empty] | NA |  | Technical difficulty to protect the java package. |  | FLASH Download Tool | UTIL_FLASH_DOWNLOAD | FLASH Download Tool |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package.
HHS Scans | UTIL_HHS_SCANS | troubleshooting | Utility to run a predefined set of scans for regulatory compliance |  |  | X | [6 empty] | X |  |  | 3 | can save 1 hour bc 30 scans and protocoles and otherwise needs to do it manually in class A ( today we have a doc in class A explaining how to do manually) |  |  | HHS Scans | UTIL_HHS_SCANS | HHS Scans |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
HHS Report | UTIL_HHS_REPORT | troubleshooting | Utility to display a complete report of the results of the HHS scans. |  |  | X | [6 empty] | X |  |  | 3 | get report - in class A has to write it paper and pen |  |  | HHS Report | UTIL_HHS_REPORT | HHS Report |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
Auto CT # Adjust (A&C are same) | CAL_WATER | troubleshooting+tube replacement | starts the CT# adjust utility |  |  | X |  |  |  | X | [5 empty] | 4 |  |  |  | Auto CT # Adjust | UTIL_AUTO_CT_NUMBER_ADJUST | Auto CT # Adjust |  |  | X | X |  |  |  | X |  |  | X |  |  | 4
Heat Soak and Seasoning (A&C are same) | UTIL_HEAT_SOAK_SEASONING | troubleshooting+tube replacement | Tube utility |  |  | X |  |  |  | X | [5 empty] | 3 |  |  |  | Heat Soak and Seasoning | UTIL_HEAT_SOAK_SEASONING | Heat Soak and Seasoning |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
CT Mobile | UTIL_CT_MOBILE | clinical use+troubleshooting+tube replacement | Configure for mobile environment (multi location) |  | X |  |  |  | X | [6 empty] | NA |  | should be A1, same as kl16 |  | CT Mobile | UTIL_CT_MOBILE | CT Mobile |  |  | X |  |  |  | X
Verify Security | UTIL_VERIFY_SECURITY | troubleshooting | check status of security key |  | X |  |  |  | X | [6 empty] | NA |  | should be A1, same as kl16 |  | Verify Security | UTIL_VERIFY_SECURITY | Verify Security |  |  | X |  |  |  | X
Editor | UTIL_EDITOR | clinical use+troubleshooting | basic editor window "gedit" |  | X | [6 empty] | X |  |  |  | 1 |  | Linux built-in or open source tool, no need to protect it. |  | Editor | UTIL_EDITOR | Editor |  |  | X | [6 empty] | X |  |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Calculator | UTIL_CALCULATOR | clinical use+troubleshooting | calculator |  | X | [6 empty] | X |  |  |  | 1 |  | Linux built-in or open source tool, no need to protect it. |  | Calculator | UTIL_CALCULATOR | Calculator |  |  | X | [6 empty] | X |  |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Calendar No tool for Cj | UTIL_CALENDAR | clinical use+troubleshooting | calendar display |  |  | X | [6 empty] | X | [6 empty] | Calendar | UTIL_CALENDAR | Calendar |  |  |  | X | [6 empty] | X |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Shell | UTIL_SHELL | clinical use+troubleshooting | opens unix shell "xterm" |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Shell | UTIL_SHELL | Shell |  |  | X |  |  |  | X
Change Display IP | UTIL_CHANGE_DISPLAY_IP | clinical use+troubleshooting | video redirect for insite |  | X |  |  |  | X | [6 empty] | NA |  | This function writes the IP to a temp file. No protection to the temp file. No need to protect the web page. |  | Change Display IP | UTIL_CHANGE_DISPLAY_IP | Change Display IP |  |  | X | [6 empty] | X |  |  |  | 1 | This function writes the IP to a temp file. No protection to the temp file. No need to protect the web page.
Software Download Update Management |  | troubleshooting |  |  |  | X | [8 empty] | X |  | Name is Update Management at Cj.
Query Updates | UTIL_QUERY_UPDATES | troubleshooting | Software download |  |  | X | [8 empty] | X | 3 |  |  |  | Query Updates | UTIL_QUERY_UPDATES | Query Updates |  |  |  | X | [6 empty] | X |  |  | 3
Download Updates | UTIL_DOWNLOAD_UPDATES | troubleshooting | Software download |  |  | X | [8 empty] | X | 3 |  |  |  | Download Updates | UTIL_DOWNLOAD_UPDATES | Download Updates |  |  |  | X | [6 empty] | X |  |  | 3
Update Status | UTIL_UPDATE_STATUS | troubleshooting | Software download |  |  | X | [8 empty] | X | 3 |  |  |  | Update Status | UTIL_UPDATE_STATUS | Update Status |  |  |  | X | [6 empty] | X |  |  | 3
Update Save | UTIL_UPDATE_SAVE |  |  |  |  | X | [8 empty] | X | 3 |  |  |  | Update Save | UTIL_UPDATE_SAVE | Update Save |  |  |  | x | [6 empty] | X |  |  | 3
Update Restore | UTIL_UPDATE_RESTORE | troubleshooting | Software download |  |  | X | [8 empty] | X | 3 |  |  |  | Update Restore | UTIL_UPDATE_RESTORE | Update Restore |  |  |  | X | [6 empty] | X |  |  | 3
Auto Download/Install Status | UTIL_ENABLE_DISABLE_AUTO_INSTALL | troubleshooting | Software download |  |  | X | [8 empty] | X | 3 |  |  |  | Enable/Disable Auto Install | UTIL_ENABLE_DISABLE_AUTO_INSTALL | Enable/Disble Auto Install |  |  |  | X | [6 empty] | X |  |  | 3
Enable/Disble Auto Download | UTIL_ENABLE_DISABLE_AUTO_DOWNLOAD | troubleshooting | Software download |  |  | X | [8 empty] | X | 3 |  |  |  | Enable/Disble Auto Download | UTIL_ENABLE_DISABLE_AUTO_DOWNLOAD | Enable/Disble Auto Download |  |  |  | X | [6 empty] | X |  |  | 3
Enable/Disble Auto Install | UTIL_AUTO_DOWNLOAD_INSTALL_STATUS | troubleshooting | Software download |  |  | X | [8 empty] | X | 3 |  |  |  | Auto Download/Install Status | UTIL_AUTO_DOWNLOAD_INSTALL_STATUS | Auto Download/Install Status |  |  |  | X | [6 empty] | X |  |  | 3
System State | UTIL_SYSTEM_STATE | clinical use+troubleshooting | Software config save/restore utility |  | X |  |  |  | X | [10 empty] | System State | UTIL_SYSTEM_STATE | System State |  |  | X |  |  |  | X
System State - USB | UTIL_SYSTEM_STATE_USB |  |  |  | X |  |  |  | X | [6 empty] | NA |  |  |  | System State - USB | UTIL_SYSTEM_STATE_USB |  |  |  | X |  |  |  | X
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | clinical use+troubleshooting | HIPAA setup utility, setting up user accounts | [5 empty] | X | [6 empty] | NA |  |  |  | EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | EA3 Admin Browser |  |  | X |  |  |  | X
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | clinical use+troubleshooting | apps shutdown |  | X |  |  |  | X | [6 empty] | NA |  |  |  | Application Shutdown | UTIL_APPLICATION_SHUTDOWN | Application Shutdown |  |  | X |  |  |  | X
Replacement (CSD menu) | [19 empty] | Replacement (CSD menu)
New Tube Config | REP_NEW_TUBE_CONFIG | troubleshooting+tube replacement | Documenting new tube ID |  | X | [5 empty] | X |  |  |  |  | 3 |  |  |  | New Tube Config | REP_NEW_TUBE_CONFIG | New Tube Config |  |  | X | [5 empty] | X |  |  |  |  | 3
FRDM Wizard | REP_FRDM_WIZARD | troubleshooting | Detector module replacement process utility |  | X |  |  |  |  | X | [5 empty] | 1 |  |  |  | FRDM Wizard | REP_FRDM_WIZARD | FRDM Wizard |  |  | X |  |  |  |  | X | [5 empty] | TBD | This tool is not found for KL16. TBD for CJ.
Product Locator | N/A |  |  |  | X |  |  |  | X | [6 empty] | NA |  | folder
Login | N/A |  |  |  | X |  |  |  | X | [6 empty] | 1
History of Change | N/A |  |  |  | X |  |  |  | X | [6 empty] | 1
Create Baseline | N/A |  |  |  | X |  |  |  | X | [6 empty] | 1
Display Baseline | N/A |  |  |  | X |  |  |  | X | [6 empty] | 1
Suite Information | N/A |  |  |  | X |  |  |  | X | [6 empty] | 1
GIB Information | N/A |  |  |  | X |  |  |  | X | [6 empty] | 1
Current Data | N/A |  |  |  | X |  |  |  | X | [6 empty] | 1
PM (CSD menu) | [19 empty] | PM (CSD menu)
Service Notepad | PM_SERVICE_NOTEPAD | troubleshooting | notepad utility for commenting into session log |  |  | X | [6 empty] | X |  |  | 3 |  |  |  | Service Notepad | PM_SERVICE_NOTEPAD | Service Notepad |  |  |  | X | [6 empty] | X |  |  | 3
Session Log | PM_SESSION_LOG | troubleshooting | session log viewer |  |  | X | [6 empty] | X |  |  | 2 |  | Technical difficulty to protect the java package. |  | Session Log | PM_SESSION_LOG | Session Log |  |  |  | X | [6 empty] | X |  |  | 2 | Technical difficulty to protect the java package.
System Health Page | PM_SYSTEM_HEALTH_PAGE | troubleshooting | Display system component status and break down to likely root cause when status is not "green". |  |  | X | [6 empty] | X |  |  | 3 |  |  |  | System Health Page | PM_SYSTEM_HEALTH_PAGE | System Health Page |  |  |  | X | [6 empty] | X |  |  | 3
Note on W14: Tao, Zhou (GE Healthcare): A Linux built-in command, no need to protect
Note on Y16: 212007413: Check A,C, and M functionality. Question if there are separate functions for each class
Note on W17: Tao, Zhou (GE Healthcare): A tool on Generator Tool JEDI which is enabled for Class C leve. Now used a special token
Note on A18: check if required in replacement or install procedure take a serie of 8/10 scans
Note on Y18: check if required in replacement or install procedure take a serie of 8/10 scans
Note on A19: eng use it.used for troubleshooting.
Note on Y19: eng use it.used for troubleshooting.
Note on A21: develop a reduce set of class A
Note on T21: class A runs the test and give errors in syslog==>give away a lot -
Note on Y21: develop a reduce set of class A QIF only for Class C
Note on W22: Tao, Zhou (GE Healthcare): Need classification on this item. This tool was enabled in DAS tools for Class C.
Note on T24: One of the most powerfull tools. For engineeirng protocle name, tracking…you can make it do the coolimeter test, the axial test.not explained in process - should be in class M.there is no procedure for it.
Note on W28: 212007413: Check to see if Kunlun16 with NIO has RAC diags. Check RT16
Note on A34: Dave: should it be moved to C ? As by replacing the TGP…or using a miror to see the LEDs on the TGP board. ( table gantry processor)
Note on Y34: Dave: should it be moved to C ? As by replacing the TGP…or using a miror to see the LEDs on the TGP board. ( table gantry processor)
Note on T37: phantom need to be relabeleed, sw change with sw doc, service procedure & reference to procedure sw( 2 sw : 1 matlab and 1 accs version, old C code)+doc = 1man month depends on sw release
Note on A55: all calibration
Note on Y55: all calibration
Note on A60: Mkg would like to move it to class C
Note on T60: Filliat, Claire (GE Healthcare): need added work ot move to C : sw and service doc
Note on Y60: Mkg would like to move it to class C
Note on Y67: required for tube + tank
Note on W92: Tao, Zhou (GE Healthcare): No A level for HHS Scans and HHS Report?
Note on W118: Tao, Zhou (GE Healthcare): Tool not found on KL16

## Optima CT520 Toolslist-TangX
Optima CT520 phase1.5 Service tools reclassification
System tools/utilities/diags | Tool License Token | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | protection level
Error Logs (CSD menu)
System Error Log | ERROR_SYSTEM_LOG |  |  |  |  | X |  |  | 2
Software History Log | ERROR_SOFTWARE_HISTORY_LOGS |  |  |  |  | X |  |  | 3
Storelog | ERROR_STORE_LOG |  | X | [5 empty] | 3
System Browser (All features ) | ERROR_SYSTEM_BROWSER |  |  |  | X |  |  |  | 1
Tube Usage | ERROR_TUBE_USAGE |  |  |  | X |  |  |  | 1
LogViewer | ERROR_LOG_VIEWER |  |  |  | X |  |  |  | 3
StartGFV | ERROR_START_GFV |  |  |  |  | X |  |  | 1
Diagnostics (CSD menu)
LAN Diagnostics | DIAG_LAN |  |  |  |  | X |  |  | 3
CPU Processes | DIAG_CPU_PROCESSES |  |  |  |  | X |  |  | 1
kV & mA (X-ray) | DIAG_KV_MA |  | X | [5 empty] | 2
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI |  | X |  |  | X |  | X | 4
Generator Tool – JEDI Filament Cal | CAL_GENERATOR_FILAMENT |  | X | [5 empty] | 4
Generator Tool – JEDI HHS Scans | UTIL_HHS_SCANS |  |  |  |  | X |  |  | 4
Generator Tool – JEDI HHS Report | UTIL_HHS_REPORT |  |  |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics Inverter Short Circuit | DIAG_GENERATOR_TOOL_JEDI |  | X |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics No Load KV | DIAG_GENERATOR_TOOL_JEDI |  | X |  |  | X |  |  | 4
Generator Tool – JEDI KV Diagnostics No Load KV(Rotation) | [7 empty] | X | 4
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE |  |  |  | X |  |  |  | 4
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL |  |  |  | X |  |  |  | 2
Axial Functional | DIAG_AXIAL_FUNCTIONAL |  |  |  |  | X |  |  | 3
DAS Tools | DIAG_DAS_TOOLS |  | X |  |  | X |  |  | 4
Scan Data Path | DIAG_SCAN_DATA_PATH |  |  |  | X |  |  |  | 2
Diagnostic Data Collection (DDC) | DIAG_DATA_COLLECTION |  |  |  | X |  |  |  | 4
DIP Diagnostics | DIAG_DIP |  |  |  | X |  |  |  | 1
High Speed Disk Test | DIAG_HIGH_SPEED_DISK |  |  |  |  | X |  |  | 1
Recon Data Path | DIAG_RECON_DATA_PATH |  |  |  | X |  |  |  | 2
scsistat OC | DIAG_SCSISTAT_OC |  |  |  |  | X |  |  | 1
rhapSnap | DIAG_RHAP_SNAP |  | X | [5 empty] | 1
GSCB Keyboard Test | DIAG_GSCB_KEYBOARD |  |  |  |  | X |  |  | 2
RTS Viewer (Basic) | DIAG_RTS_VIEWER |  |  |  | X |  |  |  | 3
RTS Viewer (Advanced) | DIAG_RTS_VIEWER_C |  |  |  |  | X |  |  | 4
RTS Log Viewer |  | X | [6 empty] | NA | marmalade
System Toubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | X | [6 empty] | NA
Console Diagnostics | DIAG_CONSOLE |  |  |  |  | X |  |  | 1
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | [6 empty] | X | 3
Smart kV Tool | DIAG_SMART_KV_TOOL | [5 empty] | X |  | 4
Image Quality (CSD menu)
Image Analysis | IQ_IMAGE_ANALYSIS | X | [6 empty] | NA
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | X | [6 empty] | NA
Scan Analysis | IQ_SCAN_ANALYSIS |  |  |  | X |  |  |  | 2
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION |  | X | [5 empty] | 3
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES | X | [6 empty] | NA
Calibration (CSD menu)
BOW Alignment |  |  | x | [5 empty] | 4
Generator Filament Calibration | CAL_GENERATOR_FILAMENT |  | X | [5 empty] | 4
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION |  | X | [5 empty] | 4
CBF and SAG Alignment |  |  | x | [5 empty] | 4
Collimator Cal | CAL_COLLIMATOR |  | X | [5 empty] | 4
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT |  |  | X |  |  |  |  | 4
ISO Alignment | CAL_ISO_ALIGNMENT |  |  | X |  |  |  |  | 4
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION |  | X | [5 empty] | 4
Default Patient Positioning | CAL_DEFAULT_PATIENT_POSITIONING | X | [6 empty] | NA
POR Alignment | CAL_POR_ALIGNMENT |  | x | [5 empty] | 3
Air cal(if no cals loaded) | NA | x | [6 empty] | 4
Detailed Calibration | CAL_DETAILED |  | X | [5 empty] | 4
Configuration (CSD menu)
Config Tracker | CONFIG_TRACKER |  | X | [5 empty] | 1
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | X | [6 empty] | NA
Install Options | CONFIG_INSTALL_OPTIONS | X | [6 empty] | NA
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID |  | X | [5 empty] | 3
Install Camera | CONFIG_INSTALL_CAMERA | X | [6 empty] | NA
Verify Options | CONFIG_VERIFY_OPTIONS | X | [6 empty] | NA
Tube Install Certification with Non-GE tube | CONFIG_TUBE_INSTALL_CERTIFICATION |  |  | X |  |  |  |  | 3
Tube Install Certification with GE Tube | CONFIG_TUBE_INSTALL_CERTIFICATION | [6 empty] | X | 3
Configure PNF | CONFIG_PNF | X | [6 empty] | NA
Configure Enhanced PNF | CONFIG_ENH_PNF | X | [6 empty] | NA | marmalade
Configure EA3 | CONFIG_EA3 | X | [6 empty] | NA
Configure EAT | CONFIG_EAT | X | [6 empty] | NA
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | X | [6 empty] | NA
Configure AW Server Client | NA | X | [6 empty] | NA
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | X | [6 empty] | NA
Component Tracker | NA
Axial | CONFIG_COMPONENT_TRACKER_AXIAL |  |  |  |  | X |  |  | 1
Console | CONFIG_COMPONENT_TRACKER_CONSOLE |  |  |  |  | X |  |  | 1
Detector | CONFIG_COMPONENT_TRACKER_DETECTOR |  |  |  |  | X |  |  | 1
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE |  |  |  |  | X |  |  | 1
History | CONFIG_COMPONENT_TRACKER_HISTORY |  |  |  |  | X |  |  | 1
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA |  |  |  |  | X |  |  | 1
Patient Information on Xtream Display | N/A | X | [6 empty] | NA
InSite Status | CONFIG_INSITE_STATUS | X | [6 empty] | NA
Configure InSite | CONFIG_INSITE |  |  |  |  | X |  |  | 3
Configure PRODIAGS | CONFIG_PRODIAGS | [6 empty] | X | 2
Configure Software Download | CONFIG_SW_DOWNLOAD |  | X | [5 empty] | 2
Configure AWS on VM | None
AWS Client Install/Uninstall | CONFIG_AWVS_CLIENT_INSTALL | X | [6 empty] | NA
Configure AWS Applications | CONFIG_AWVS_APPLICATIONS | X | [6 empty] | NA
Launch AWS Service Tool | CONFIG_LAUNCH_AWVS_SERVICE | X | [6 empty] | NA
Gantry Table User Preference | CONFIG_GANTRY_TABLE_PREFS | X | [6 empty] | NA
Configure HIPAA | CONFIG_HIPAA | X | [6 empty] | NA
Utilities (CSD menu)
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | X | [6 empty] | NA
HHS Scans(Class A manual exist method) | UTIL_HHS_SCANS |  |  |  |  | X |  |  | 3
HHS Report | UTIL_HHS_REPORT |  |  |  |  | X |  |  | 3
Manual CT # Adjust | UTIL_MANUAL_CT_ADJUST |  | X | [5 empty] | 4
Auto CT # Adjust | CAL_WATER |  | X | [5 empty] | 4
Heat Soak and Seasoning (A&C are same) | UTIL_HEAT_SOAK_SEASONING |  | X | [5 empty] | 3
CT Mobile | UTIL_CT_MOBILE | X | [6 empty] | NA
Verify Security | UTIL_VERIFY_SECURITY | X | [6 empty] | NA
Editor | UTIL_EDITOR |  |  |  | X |  |  |  | 1
Calculator | UTIL_CALCULATOR |  |  |  | X |  |  |  | 1
Shell | UTIL_SHELL | X | [6 empty] | NA
Change Display IP | UTIL_CHANGE_DISPLAY_IP | X | [6 empty] | NA
System State（removed in TangX） | UTIL_SYSTEM_STATE | X
System State - USB | UTIL_SYSTEM_STATE_USB | X | [6 empty] | NA
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | X | [6 empty] | NA
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | X | [6 empty] | NA
TVA | UTIL_TVA | X | [6 empty] | NA
Install Software Patch | UTIL_INSTALL_SW_PATCH | [6 empty] | X | 1
Security Center | UTIL_SECURITY_CENTER | X | [6 empty] | NA
Reset Admin User Password | UTIL_RST_ADM_PASSWD | [6 empty] | X | NA | marmalade
Reset Service User Password | UTIL_RST_SRV_PASSWD | [6 empty] | X | NA | marmalade
Audit Log Viewer | UTIL_AUDIT_LOG_VIEWER | X | [6 empty] | NA
Replacement (CSD menu)
New Tube Config | REP_NEW_TUBE_CONFIG |  |  | X |  |  |  |  | 3
FRDM Wizard | REP_FRDM_WIZARD |  | X | [5 empty] | 1
Product Locator | N/A | X | [6 empty] | NA
PM (CSD menu)
Service Notepad | PM_SERVICE_NOTEPAD |  |  |  |  | X |  |  | 3
Session Log | PM_SESSION_LOG |  |  |  |  | X |  |  | 2
System Health Page | PM_SYSTEM_HEALTH_PAGE |  |  |  |  | X |  |  | 3
Scan Utilities
Scan Utilties - Detailed Carlibration | CAL_DETAILED |  | x | [5 empty] | 4
Scan Utilties - Center Phantom | CAL_CENTER_PHANTOM | X | [6 empty] | 4
Scan Utilties - Auto CT Number Adjust | CAL_WATER |  | x | [5 empty] | 4
Scan Utilties - Collimator Cal | CAL_COLLIMATOR |  | x | [5 empty] | 4
Scan Utilties - Air Calibration | CAL_AIR | X | [6 empty] | 4
Scan Utilties - Spectral Cal | CAL_SPECTRAL |  |  | X |  |  |  |  | 4

## KL16 Tools&Utilities List (hidden)
KL16 Service tools reclassification
view Service Document |  |  |  | Current IB Classification |  |  |  | New IB Classification - Cj
System tools/utilities/diags (As they exist on KL16 today, very similar to older CT IB systems as well) | Tool License Token | CJ tools name | Tool overview |  | Class A | Class C | Class M |  | Class A 1 ( Clinical use) | Class A2 ( Troubleshooting) | Class A3 ( tube replacement) | Class C1 (was in A in the past) | Class C 2 ( wasin class C) | Class C3 ( was in M) | Class M | Protection Level | Comments | Impact of reclassification to Indepedant Service Provider (Class A moved to Class C)
Error Logs (CSD menu)
System Error Log | ERROR_SYSTEM_LOG | System Error Log |  |  |  | X | [6 empty] | X |  |  | 2 | No protection to log files anyway. And it's using a viewlog command which may be widely used.
Software History Logs | ERROR_SOFTWARE_HISTORY_LOGS | Software History Log |  |  |  | X | [6 empty] | X |  |  | 3
Store Log | ERROR_STORE_LOG | Storelog |  |  | X |  |  |  |  | X | [5 empty] | 3
System Browser(All feature) | ERROR_SYSTEM_BROWSER | System Browser (All features ) |  |  | X | [6 empty] | X |  |  |  | 1 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
Tube Usage | ERROR_TUBE_USAGE | Tube Usage |  |  | X | [6 empty] | X |  |  |  | 1 | No protection to log files anyway. Technical difficulty due to it's using CGI scripts with web param.
LogViewer | ERROR_LOG_VIEWER | LogViewer |  |  | X | [6 empty] | X |  |  |  | 3
StartGFV | ERROR_START_GFV | StartGFV |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty: it's only a tar package installed to console.
Diagnostics (CSD menu)
LAN Diagnositics | DIAG_LAN | LAN Diagnostics |  |  |  | X | [6 empty] | X |  |  | 3
CPU Processes | DIAG_CPU_PROCESSES | CPU Processes |  |  |  | X | [6 empty] | X |  |  | 1 | This is a Linux built-in command. No need to protect it.
kV & mA (X-ray) | DIAG_KV_MA | kV & mA (X-ray) |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Generator Tool - JEDI | DIAG_GENERATOR_TOOL_JEDI | Generator Tool - JEDI |  |  | X | X | X |  |  | X |  |  | X |  | X | 3 | Need to include KVMA in license
Generator Tool JEDI - Focal Spot | DIAG_FOCAL_SPOT | [14 empty] | 4 | Need classification.
Collimator Aperture Test | DIAG_COLLIMATOR_APERTURE | Collimator Aperture Test |  |  | X | [6 empty] | X |  |  |  | 3
Collimator Functional | DIAG_COLLIMATOR_FUNCTIONAL | Collimator Functional |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Axial Functional | DIAG_AXIAL_FUNCTIONAL | Axial Functional |  |  |  | X | [6 empty] | X |  |  | 3
DAS Tools | DIAG_DAS_TOOLS | DAS Tools |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
DAS Tools - QIF Vis | DIAG_DAS_TOOLS_QIF_VIS |  |  |  |  | X | [9 empty] | 4 | Need classification.
Scan Data Path | DIAG_SCAN_DATA_PATH | Scan Data Path |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
Diagnostic Data Collection (DDC) | DIAG_DATA_COLLECTION | Diagnatoc Data Collection (DDC) |  |  | X | [6 empty] | X |  |  |  | 4 | DDC tool has problem with L3 protection, so changed to Level 4.
DIP Diagnostics | DIAG_DIP | VDIP Diagnostics |  |  | X | [6 empty] | X |  |  |  | 3
High Speed Disk Test | DIAG_HIGH_SPEED_DISK | High Speed Disk Test |  |  |  | X | [6 empty] | X |  |  | 3
Recon Data Path | DIAG_RECON_DATA_PATH | Recon Data Path |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
RAC Diagnostics | DIAG_RAC | VRAC Diagnostics | [13 empty] | 3
scsistat OC | DIAG_SCSISTAT_OC | scsistat OC |  |  |  | X | [6 empty] | X |  |  | 1 | This command is used by software startup, cannot protect it.
rhapSnap | DIAG_RHAP_SNAP | rhapSnap |  |  |  | X | [6 empty] | X |  |  | 3
GSCB Keyboard Test | DIAG_GSCB_KEYBOARD | SCIM Keyboard Test |  |  |  | X | [6 empty] | X |  |  | 2 | Technical difficulty to protect the java package because it's shared by many programs.
RTS viewer | DIAG_RTS_VIEWER | RTS Viewer A |  |  | X | [6 empty] | X |  |  |  | 3
RTS viewer C | DIAG_RTS_VIEWER_C | RTS Viewer C |  |  |  | X | [6 empty] | X |  |  | 4
System Troubleshooting Tool | DIAG_SYSTEM_TROUBLESHOOTING | System Toubleshooting Tool |  |  | X |  |  |  | X
Console Diagnostics | DIAG_CONSOLE | Console Diagnostics |  |  |  | X | [6 empty] | X |  |  | 3 | Need to include DIP Diag in license
System Health Monitor | DIAG_SYSTEM_HEALTH_MONITOR | System Health Monitor |  |  |  |  | X | [7 empty] | X | 3 | Whether we can wrap /vobs/srv/Tools/shm/UI/shmUI/dist/shmui.jar too?
Smart kV Tool | DIAG_SMART_KV_TOOL | Smart kV Tool |  |  |  |  | X | [6 empty] | X |  | 3
Image Quality (CSD menu)
Image Analysis | IQ_IMAGE_ANALYSIS | Image Analysis |  |  | X |  |  |  | X
Install SMPTE Image | IQ_INSTALL_SMPTE_IMAGE | Install SMPTE Image |  |  | X |  |  |  | X
Scan Analysis | IQ_SCAN_ANALYSIS | Scan Analysis |  |  | X | [6 empty] | X |  |  |  | 2 | Technical difficulty to protect the java package.
Install Image Orientation Imgs | IQ_INSTALL_IMAGE_ORIENTATION | Install Image Orientation Imgs |  |  | X |  |  |  |  | X | [5 empty] | NA | No protection to the images.
Install 21CFR Images | IQ_INSTALL_21CFR_IMAGES |  |  |  | X |  |  |  | X | [6 empty] | 3 | No protection to the images.
Calibration (CSD menu)
Generator Filament Calibration | CAL_GENERATOR_FILAMENT | Generator Filament Calibration |  |  | X |  |  |  |  | X | [5 empty] | 3
Gantry Rotation Characterization | CAL_GANTRY_ROTATION_CHARACTERIZATION | Gantry Rotation Characterization |  |  | X |  |  |  |  | X | [5 empty] | 3
Collimator Cal | CAL_COLLIMATOR | Collimator Cal |  |  | X |  |  |  |  | X | [5 empty] | 4
Hot ISO Alignment | CAL_HOT_ISO_ALIGNMENT | Hot ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 3
ISO Alignment | CAL_ISO_ALIGNMENT | ISO Alignment |  |  | X | [5 empty] | X |  |  |  |  | 3
Mechanical Characterization | CAL_MECHANICAL_CHARACTERIZATION | Mechanical Characterization |  |  | X |  |  |  |  | X | [5 empty] | 3
Default Patient Positioning | CAL_DEFAULT_PATIENT_POSITIONING | Default Patient Positioning |  |  | X |  |  |  | X
Detailed Calibration | CAL_DETAILED | Detailed Calibration |  |  | X | [5 empty] | X |  |  |  |  | 4
BOW Alignment | CAL_BOW_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
POR Alignment | CAL_POR_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
CBF and SAG Alignment | CAL_CBF_SAG_ALIGNMENT |  |  |  | X |  |  |  |  | X | [5 empty] | 3
| CAL_Z_ALIGN | Z Align |  |  | X |  |  |  |  | X | [5 empty] | TBD | Tool not used by KL16, TBD for CJ.
| CAL_SMART_TUBE_ALIGNMENT | Smart Tube Alignment |  |  |  |  | X | [6 empty] | X |  | TBD | Tool not used by KL16, TBD for CJ.
| CAL_GANTRY_SERVICE_BALANCE | Gantry Service Balance |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package.
Scan Utilties - Center Phantom | CAL_CENTER_PHANTOM |  |  |  | x |  |  |  | x | [6 empty] | 4 | This is a button on Scan utilities.
Scan Utilties - Z Slope Cal | CAL_Z_SLOPE |  |  |  | x | [5 empty] | X |  |  |  |  | 4 | This is a button on Scan utilities.
Scan Utilties - DAS Gain Cal | CAL_DAS_GAIN |  |  |  | x | [5 empty] | x |  |  |  |  | 4 | This is a button on Scan utilities.
Scan Utilties - Air Cal | CAL_AIR |  |  |  | x |  |  |  | x | [6 empty] | 4 | This is a button on Scan utilities.
Scan Utilties - Spectral Cal | CAL_SPECTRAL |  |  |  | x | [5 empty] | x |  |  |  |  | 4 | This is a button on Scan utilities.
Configuration (CSD menu) |  | Transfer Function Calibration
Config Tracker | CONFIG_TRACKER | Config Tracker |  |  | X |  |  |  |  | X | [5 empty] | 1 | This tool is just a display of config info (which is generated after software LFC). Currently we cannot protect it.
Injector: Pressure Limit Unit | CONFIG_INJECTOR_PRESSURE_LIMIT_UNIT | Injector: Pressure Limit Unit |  |  | X |  |  |  | X
Install Options | CONFIG_INSTALL_OPTIONS | Install Options |  |  | X |  |  |  | X
Product SW Revs | CONFIG_PRODUCT_SW_REVS | Product SW Revs |  |  | X |  |  |  |  | X | [5 empty] | 1 | This command is used by software startup, cannot protect it.
Enter Detector ID | CONFIG_ENTER_DETECTOR_ID | Enter Detector ID |  |  | X |  |  |  |  | X | [5 empty] | 3
Install Camera | CONFIG_INSTALL_CAMERA | Install Camera |  |  | X |  |  |  | X
Verify Options | CONFIG_VERIFY_OPTIONS | Verify Options |  |  | X |  |  |  | X
Tube Install Certification | CONFIG_TUBE_INSTALL_CERTIFICATION | Tube Install Certification |  |  | X |  | X |  |  |  | X |  |  |  | X | 2
Configure PNF | CONFIG_PNF | Configure PNF |  |  | X |  |  |  | X
Configure EA3 | CONFIG_EA3 | Configure EA3 |  |  | X |  |  |  | X
Configure EAT | CONFIG_EAT | Configure EAT |  |  | X |  |  |  | X
Configure AW DirectConnect | CONFIG_AW_DIRECT_CONNECT | Configure AW DirectConnect |  |  | X |  |  |  | X
Install Gating Monitor | CONFIG_INSTALL_GATING_MONITOR | Install Gating Monitor |  |  | X |  |  |  | X
Component Tracker |  | Component Tracker
Console | CONFIG_COMPONENT_TRACKER_CONSOLE | Console |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
Software | CONFIG_COMPONENT_TRACKER_SOFTWARE | Software |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
History | CONFIG_COMPONENT_TRACKER_HISTORY | History |  |  |  | X | [6 empty] | X |  |  | 1 | Technical difficulty to protect such payload in the form of servlet.
Refresh Console Data | CONFIG_COMPONENT_TRACKER_REFRESH_CONSOLE_DATA | Refresh Console Data |  |  |  | X | [6 empty] | X |  |  | 3
| CONFIG_COMPONENT_TRACKER_DETECTOR | Detector | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
|  | Configure AW Server Client | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
|  | Patient Information on Xtream Display | [13 empty] | TBD | Tool not used by KL16, TBD for CJ.
Utilities (CSD menu)
FLASH Download Tool | UTIL_FLASH_DOWNLOAD | FLASH Download Tool |  |  | X |  |  |  |  | X | [5 empty] | 2 | Technical difficulty to protect the java package.
HHS Scans | UTIL_HHS_SCANS | HHS Scans |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
HHS Report | UTIL_HHS_REPORT | HHS Report |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
Auto CT # Adjust | UTIL_AUTO_CT_NUMBER_ADJUST | Auto CT # Adjust |  |  | X | X |  |  |  | X |  |  | X |  |  | 4
Heat Soak and Seasoning | UTIL_HEAT_SOAK_SEASONING | Heat Soak and Seasoning |  |  | X | X |  |  |  | X |  |  | X |  |  | 3
CT Mobile | UTIL_CT_MOBILE | CT Mobile |  |  | X |  |  |  | X
Verify Security | UTIL_VERIFY_SECURITY | Verify Security |  |  | X |  |  |  | X
Editor | UTIL_EDITOR | Editor |  |  | X | [6 empty] | X |  |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Calculator | UTIL_CALCULATOR | Calculator |  |  | X | [6 empty] | X |  |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Calendar | UTIL_CALENDAR | Calendar |  |  |  | X | [6 empty] | X |  |  | 1 | Linux built-in or open source tool, no need to protect it.
Shell | UTIL_SHELL | Shell |  |  | X |  |  |  | X
Change Display IP | UTIL_CHANGE_DISPLAY_IP | Change Display IP |  |  | X | [6 empty] | X |  |  |  | 1 | This function writes the IP to a temp file. No protection to the temp file. No need to protect the web page. It will be update to be identical with CJ later
Query Updates | UTIL_QUERY_UPDATES | Query Updates |  |  |  | X | [6 empty] | X |  |  | 3
Download Updates | UTIL_DOWNLOAD_UPDATES | Download Updates |  |  |  | X | [6 empty] | X |  |  | 3
Update Status | UTIL_UPDATE_STATUS | Update Status |  |  |  | X | [6 empty] | X |  |  | 3
Update Save | UTIL_UPDATE_SAVE | Update Save |  |  |  | x | [6 empty] | X |  |  | 3
Update Restore | UTIL_UPDATE_RESTORE | Update Restore |  |  |  | X | [6 empty] | X |  |  | 3
Enable/Disable Auto Install | UTIL_ENABLE_DISABLE_AUTO_INSTALL | Enable/Disble Auto Install |  |  |  | X | [6 empty] | X |  |  | 3
Auto Download/Install Status | UTIL_AUTO_DOWNLOAD_INSTALL_STATUS | Auto Download/Install Status |  |  |  | X | [6 empty] | X |  |  | 3
Enable/Disble Auto Download | UTIL_ENABLE_DISABLE_AUTO_DOWNLOAD | Enable/Disble Auto Download |  |  |  | X | [6 empty] | X |  |  | 3
System State | UTIL_SYSTEM_STATE | System State |  |  | X |  |  |  | X
System State - USB | UTIL_SYSTEM_STATE_USB |  |  |  | X |  |  |  | X
EA3 Admin Browser | UTIL_EA3_ADMIN_BROWSER | EA3 Admin Browser |  |  | X |  |  |  | X
Application Shutdown | UTIL_APPLICATION_SHUTDOWN | Application Shutdown |  |  | X |  |  |  | X
Replacement (CSD menu)
New Tube Config | REP_NEW_TUBE_CONFIG | New Tube Config |  |  | X | [5 empty] | X |  |  |  |  | 3
FRDM Wizard | REP_FRDM_WIZARD | FRDM Wizard |  |  | X |  |  |  |  | X | [5 empty] | 1 | This tool is not used for KL16 China. Change level to 1 for CJ.
PM (CSD menu)
Service Notepad | PM_SERVICE_NOTEPAD | Service Notepad |  |  |  | X | [6 empty] | X |  |  | 3
Session Log | PM_SESSION_LOG | Session Log |  |  |  | X | [6 empty] | X |  |  | 2 | Technical difficulty to protect the java package.
System Health Page | PM_SYSTEM_HEALTH_PAGE | System Health Page |  |  |  | X | [6 empty] | X |  |  | 3
Note on A14: Tao, Zhou (GE Healthcare): A Linux built-in command, no need to protect
Note on C16: 212007413: Check A,C, and M functionality. Question if there are separate functions for each class
Note on A17: Tao, Zhou (GE Healthcare): A tool on Generator Tool JEDI which is enabled for Class C leve. Now used a special token
Note on C18: check if required in replacement or install procedure take a serie of 8/10 scans
Note on C19: eng use it.used for troubleshooting.
Note on C21: develop a reduce set of class A QIF only for Class C
Note on A22: Tao, Zhou (GE Healthcare): Need classification on this item. This tool was enabled in DAS tools for Class C.
Note on U23: class A runs the test and give errors in syslog==>give away a lot -
Note on U25: One of the most powerfull tools. For engineeirng protocle name, tracking…you can make it do the coolimeter test, the axial test.not explained in process - should be in class M.there is no procedure for it.
Note on A28: 212007413: Check to see if Kunlun16 with NIO has RAC diags. Check RT16
Note on C34: Dave: should it be moved to C ? As by replacing the TGP…or using a miror to see the LEDs on the TGP board. ( table gantry processor)
Note on U37: phantom need to be relabeleed, sw change with sw doc, service procedure & reference to procedure sw( 2 sw : 1 matlab and 1 accs version, old C code)+doc = 1man month depends on sw release
Note on C52: all calibration
Note on C57: Mkg would like to move it to class C
Note on U58: Filliat, Claire (GE Healthcare): need added work ot move to C : sw and service doc
Note on C64: required for tube + tank
Note on A90: Tao, Zhou (GE Healthcare): No A level for HHS Scans and HHS Report?
Note on A115: Tao, Zhou (GE Healthcare): Tool not found on KL16

## Markup (hidden)
Markup | Instruction
VRAC Diagnostics | KL64已删除， Project Tang 不删除
Manual CT # Adjust | 与KL64对比 是纯新加的
Product Locator | 与KL64比需要删除的
RTS Viewer (class A) ??? | 有疑问需要讨论