## Compliance
GE Healthcare Services Quality Form - Compliance
Parent document: Optima CT680 Installation and Service Directions
Instructions
Complete this form in its entirety. Data entry fields are indicated with a Blue border. Entries must be made in all fields. When calculations are required they will be performed automatically from the data you enter. Send the completed form to Sendhhs@ge.com.
Conventions:
A data entry field
A spec limit.
A note to the FE using this form.
A major section of this document
A minor section or group of entries.
A calculated field
Customer Information
Customer name
System ID
Order number
Dispatch number
Field Engineer
Test Date
Federal / State Compliance
Scanner Information - Certified components | Model Number |  | Serial Number
Collimator
Master Control Console
Gantry
HV Generator
Auxilliary Box
Power Unit
HV Tank
HEMIT Tank
Table
Xray Tube Casing
Xray Tube Insert
System Components are installed in the locations identified by GE Healthcare drawings.
Console
Gantry
Table
GE Healthcare site drawings match State approved shielding plans.
Note: Shielding Plans are required for specific U.S.A. states of installation. The State specific requirements are: Arkansas required for new construction District of Columbia: Customer copy of the Plan is submitted. North Carolina, South Carolina, and Washington: required
Note: Enter the Customer's facility registration number on the 2579 form for the required U.S.A. states.
Test Equipment used during the installation
Note: Your personal equipment data can be used to replace portions of the following table. To use your personal data instead of the following table: 1. Create a new worksheet, titled TOOLS. 2. Import your list into the TOOLS worksheet. 3. Fill the table below with "see attached TOOLS".
DMM | Manufacturer
Model Number
Serial Number
Device is within calibration date range
Torque Wrench | Manufacturer
Model Number
Serial Number
Device is within calibration date range
From the Tool Pool:
Leakage Meter | Manufacturer
Model Number
Serial Number
Device is within calibration date range
Comments:
DOC1798981 rev 1 | GE4879 - GE Healthcare Services Quality Form |  |  | Page 1
Copyright 2016 General Electric Company. All rights reserved.

## Mechanical Install
GE Healthcare Services Quality Form - Mechanical Install
Parent document: Optima CT680 Installation and Service Directions
Table / Gantry Alignment
| Specifications |  |  |  | Test Result
| Min | Max | Units | Data
Cradle Level - X direction | Bubble centered in finder. |  | NA
Cradle Level - Z direction | Bubble centered in finder. |  | NA
Table/Gantry alignment was completed per Book 1 of the Installation Manual. | NA | NA | NA | NA
Table/Gantry alignment is within specification. | -1 | 1 | mm
Table / Gantry Anchoring
| Specifications |  |  |  | Test Result
| Min | Max | Units | Data
Gantry - Anchor type | NA | NA | NA
Gantry - Anchor supplied by | NA | NA | NA
Gantry - Anchoring conforms to the specifications | NA | NA | NA | NA
Gantry - Length of anchor extending above the securing nut is entered in the field to the right and is within the specification limits. | 24 (8-in anchor) 13 (7-in anchor) | 33 (8-in anchor) 25 (7-in anchor) | mm
Gantry - Anchor torque used is entered in the field to the right and is within the specification limits. | 8-inch (2106573): 75 (55) 7-inch (5487992-2): 54 (40) |  | N-m (lb-ft)
Table - Anchor type | NA | NA | NA
Table - Anchor supplied by | NA | NA | NA
Table - Anchoring conforms to the specifications | NA | NA | NA | NA
Table - Length of anchor extending above the securing nut is entered in the field to the right and is within the specification limits. | 24 (8-in anchor) 13 (7-in anchor) | 33 (8-in anchor) 25 (7-in anchor) | mm
Table - Anchor torque used is entered in the field to the right and is within the specification limits. | 8-inch Anchor (2106573): 75 (55) 7-inch Anchor (5487992-2): 54 (40) 140mm Anchor (2106573-3): 50 (37) |  | N-m (lb-ft)
Floor type is concrete and of the thickness is entered in the field to the right. | 102 | NA | mm
Note: If the Floor type entered is 'Other', identify the floor material and thickness in this field. Otherwise enter "NA" in this field.
Floor thickness, level, and material conforms to the specifications.
System Interconnect Cables Continuity Test
Mechanical installation procedures were followed.
All continuity tests passed.
All cables are secured within their raceways without being damaged or pinched.
Mechanical Installation Tasks
Injector installation completed.
Dollies and lean carts returned.
Gantry bearing inspection completed.
All customer ordered hardware options are installed.
Mechanical Installer Name
Mechanical Install Company
Comments:
DOC1798981 rev 1 | GE4879 - GE Healthcare Services Quality Form |  |  |  |  | Page 2
Copyright 2016 General Electric Company. All rights reserved.

## Electrical Install
GE Healthcare Services Quality Form - Electrical Install
Parent document: Optima CT680 Installation and Service Directions
System Warning Labels
Laser warning labels:
Note: The warning label installed must be written in the appropriate language for the Country the system is installed in.
Test Result
Laser warning labels are installed on the lower right corner of the Gantry (near the floor) and near each laser window opening in the Gantry covers.
Laser warning labels are in the appropriate language.
SCIM or GSCB warning labels:
Test Result
The SCIM or GSCB Xray warning indicator is attached to the SCIM.
The SCIM or GSCB Xray warning statement is in the appropriate language.
Gantry/Table warning labels |  |  |  |  | Test Result
All warning labels were installed following the System Installation Manual.
All warning labels are in the appropriate language as defined in Direction 5221102-1EN.
User Interface Language
The UI is expected to be in the language appropriate for the country in which this system is installed, according to Direction 5221102-1EN. If an authorized site representative chooses another UI language, record the following information. Otherwise enter 'NA'.
Radiology Manager (or equivalent)
Date of discussion
Language selected
Appropriate or Requested language is installed.
Alignment Lights
Internal axial lights are functional and aligned.
External axial lights are functional and aligned.
Coronal lights are functional and aligned.
External Alignment lights are functional (if present).
Table Cradle alignment light accuracy test passed.
Auto Voice
Auto Voice volume has been adjusted.
Patient speaker volume has been adjusted.
Operatore Console speaker volume has been adjusted.
Documentation
FE documentation received for all components installed.
Customer documentation received for all components installed.
Operator's manual, in the appropriate language, has been provided to the Director of Radiology (or equivalent).
System Options
Advantage 4D - installed (if present).
Cardiac Option - installed (if present).
Remote Monitor - installed (if present).
SmartStep - installed (if present).
UPS - installed (if present).
System Tests
Cable / Clearance |  |  |  |  | Test Result
Patient Touch Leakage test was performed and the results indicate a passing condition.
System network connections are complete and tested.
Table Height characterization was performed and the results indicate a passing condition.
Interference test was performed and the results indicate a passing condition. Note: The cradle extender must be installed for this test.
System Ground Chassis Leakage | Specification |  |  |  | Test Result
| Min | Max | Units | Data
Console ground wire measurement indicates a passing condition. | 0 | 5 | mA
Gantry ground wire measurement indicates a passing condition. | 0 | 5 | mA
PDU ground wire measurement indicates a passing condition. | 0 | 5 | mA
Table ground wire measurement indicate a passing condition. | 0 | 5 | mA
Boom-In-Room ground wire measurement (if present) indicates a passing condition. | 0 | 5 | mA
Tomographic Plane Indicator Accuracy | Specification |  |  |  | Test Result
| Min | Max | Units | Data
Internal landmark scan plan deviation test indicates a passing condition. | -1.0 | +1.0 | mm
External landmark scan plan deviation test indicates a passing condition. | -1.0 | +1.0 | mm
Note: If scan location S1.0 or scan location I1.0 has the darkest bar, the system meets the desired specification.
Xray On Indicators |  |  |  |  | Test Result
The SCIM or GSCB Xray On indicator illuminates during a scan. |  |  |  |  | Yes
The Gantry front cover Xray On indicator illuminates during a scan. |  |  |  |  | Yes
The Gantry rear cover Xray On indicator illuminates during a scan.
The Room Xray On indicator illuminates during a scan.
The Room door interlock stops a scan when the door is open (if present).
Special Case Installation |  |  |  | Data | Result
Is this system installed in a suite where surgical procedures are peformed?
If yes, enter CSO case number here:
Functional System Tests |  |  |  |  | Test Result
Functional System Tests was performed and the results indicate a passing condition.
Image Quality |  |  |  |  | Test Result
Image Quality test was performed and the results indicate a passing condition.
FE Completion Tasks
Scan Room | Specification |  |  | Data | Test Result
| Min | Max | Units
Scan room humidity is recorded and within specification. | 30 | 60 | %RH
Scan room temperature is recorded and within specification. | 18 | 26 | degree Centigrade
Scan room HVAC results indicate a passing condition.
Finalization |  |  |  |  | Test Result
Customer documentation was received for all components installed.
Generator recalibration was required.
Collimator recalibration was required.
System re-alignment was required.
Option DVDs are stored in the GE Service cabinet (or equivalent).
System software CDs are stored in the GE Service cabinet (or equivalent).
Comments:
DOC1798981 rev 1 | GE4879 - GE Healthcare Services Quality Form |  |  |  |  | Page 3
Copyright 2016 General Electric Company. All rights reserved.

## RevHistory
GE Healthcare Services Quality Form - Revision History
Parent document: Optima CT680 Installation and Service Directions
Revision | Date | Author | Description of Change
1 | 2016-01-08 | Zhang Ce | Initial Release.
Template Information
Document | DOC1798981
Revision | 1
Title | GE4879 - GE Healthcare Services Quality Form
Footer | Copyright 2016 General Electric Company. All rights reserved.
DOC1798981 rev 1 | GE4879 - GE Healthcare Services Quality Form |  |  | Page 4
Copyright 2016 General Electric Company. All rights reserved.