## Sheet1
EC# | Description | Location
EC1 | The system is designed to provide adequate protection against hazards caused by overflow | Service methods (5487413-8EN)->Align, Setup, Cals->Gantry->Tilt->Hydraulic Tilt Motor Fluid Check and Fill Procedure Service methods (5487413-8EN)->Replacement-> Gantry->Tube->X-Ray Tube replacement
EC2 | The system is designed to provide adequate protection against hazards caused by leakage | Service methods (5487413-8EN)->Functional Checks-> Enclosure Leakage and system chassis leakage tests
EC3 | The system is designed to provide adequate protection against hazards caused by humidity | Pre-Installation Manual 5472000-1EN chapter 9 section 1.0 Installation Manual 5472001-1EN chapter 7 section 3.0
EC4 | Covers are designed to minimise the possibility of fluid ingress. | Pre-Installation Manual 5472000-1EN chapter 9 section 2.0
EC5 | Comply with applicable standards and regulations pertaining to Radiation Control | Pre-Installation Manual 5472000-1EN chapter 10 section 1.0
EC6 | Comply with applicable standards and regulations pertaining to Xray Indication Safety | Pre-Installation Manual 5472000-1EN chapter 13 section 5.0
EC7 | Manufacturing and Service testing procedures for kV, mA, and scan time accuracy. | Service methods (5487413-8EN)->Functional Checks-> Gantry->HHS Scan
EC8 | Design uses IEC standards as a tool for safety factor definition. Manufacturing and Service procedure gives information on how to balance the gantry. | Service methods (5487413-8EN)->Align, Setup, Cals-> Gantry-> Gantry Balance Procedure
EC9 | In the event of failure of Service Process, the clearance to the stationary frame should be verified by a field service engineer before rotation and Service Methods shall provide 'Manually rotate gantry, at least one full revolution, and verify that no rubbing occurs' at Tube replacement section. | Service methods (5487413-8EN)->Replacement-> Gantry-> Tube-> X-Ray Tube Replacement
EC10 | The system tolerates the use of biologically effective cleaning agents Cleaning procedure in user and service documentation. | Service methods (5487413-8EN)->Planned Maintenance and Tasks
EC11 | Service procedure instructs proper Date and Time setting in software installation process. | Service methods (5487413-8EN)->Software-> (16BW45.x /17 BW34.x) Load From Cold
EC12 | Include periodic maintenance, daily and periodic operator checks and calibrations in service and operator manual. Objective Evidence: Service Manuals and User Instructions | Service methods (5487413-8EN)->Planned Maintenance and Tasks
EC13 | System provides a PM Schedules document along with the scanner. | Service methods (5487413-8EN)->Planned Maintenance and Tasks
EC14 | Service procedure to characterize table/gantry motion at install. | Installation Manual 5472001-1EN chapter 5 section 2.0