GE FORM – System Chassis Leakage Test


Customer Site Data
	Where Installed

	City
	State

	Customers telephone Number
	Room No.

	System ID #
	GON #

	DVM Serial #
	Cal Due Date

	Leakage Meter Manufacturer
	Model #

	Serial #
	Cal Due Date

	Installers Name:
	Company

	Test Date:


Leakage Measurements*
	
	mA

Power On
	mA

Power Off
	Test Date

	Table Ground Bus
	
	
	

	Table Ground wire
	
	
	

	PDU Ground wire
	
	
	

	Console ground wire
	
	
	

	Options
	
	
	

	Remote Monitor/ (Boom in Room)
	
	
	

	Injector
	
	
	

	Other
	
	
	

	Other
	
	
	

	Other
	
	
	

	Tester


*The measured leakage current shall not exceed 5 milli-amps in any ground wire.
Note: if there is not enough spaces to write results for any other electronic devices that must be tested use an additional form complete the Customer Site Data information and place the data in the rows designated “other”. 
Site Ground Continuity Check
	From
	To
	Measurement*

	PDU Ground  bus
	Vault Ground
	

	PDU Ground bus
	Table/Gantry raceway ground point
	

	Table/Gantry raceway ground point
	Gantry Chassis
	

	Table/Gantry raceway ground point
	Table Chassis
	

	Table/Gantry raceway ground point
	Operator Console Chassis
	

	All Display or Computing Options (if any)
	Operator Console Chassis
	

	Tester


*The measured resistance between points shall be less than 1.0 ohm All tests listed above:


□ Passed 
□ Failed
Return to: Project Manager of Installation
DOC0594445
Rev 4 (9-Jul-2014)
Page 1 of 1

GE - Confidential


