## Compliance
GE Healthcare Services Quality Form - Compliance
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Certified Component Verification Form
Copyright 2014 General Electric Company. All rights reserved.
Instructions
Complete the sections in this form that correspond to the component that was replaced. Data entry fields are indicated with a Blue border. Entries must be made in all fields for the replaced component. When calculations are required they will be performed automatically from the data you enter. Send the completed from to Sendhhs@ge.com.
Conventions:
A data entry field
A spec limit.
A note to the FE using this form.
A major section of this document
A minor section or group of entries.
A calculated field
Customer Information
Customer name
System ID
Order number
Dispatch number
Field Engineer
Test Date
State Compliance
South Carolina requires a State Approved shielding plan or Approved Post Installation Survey for the replacement of the Console or Generator. Received state approval letter prior to replacement of the console or generator.
Test Equipment used during the repair.
Note: Your personal equipment data can be used to replace the following table. To use your personal data instead of the following table: 1. Create a new worksheet, titled TOOLS. 2. Import your list into the TOOLS worksheet. 3. Fill the table below with "see attached TOOLS".
DMM | Manufacturer
Model Number
Serial Number
Calibration due date
AMP Meter or Current *Clamp | Manufacturer
Model Number
Serial Number
Calibration due date
Torque Wrench | Manufacturer
Model Number
Serial Number
Calibration due date
Oscilloscope | Manufacturer
Model Number
Serial Number
Calibration due date
From the Tool Pool:
Bleeder | Manufacturer
Model Number
Serial Number
Calibration due date
Leakage Meter | Manufacturer
Model Number
Serial Number
Calibration due date
kV Measurement & XRT Alignment Phantom | Manufacturer |  | GE Healthcare
Model Number
Serial Number
Calibration due date
Note:* If you use a Current Clamp with a Fluke meter, it should be sent for calibration at the same time as the Fluke meter. Use the DVM manufacturer, model number, serial number, and calibration date.
Certified Component | Model Number | Serial Number |  | Installation Date
Collimator
Gantry
HV Generator
Auxilliary Box
Power Unit
HV Tank
HEMIT Tank
Master Control (Console)
Table
Tube Casing
Tube Insert
Comments:
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  |  |  | Page 1
GE Confidential

## Collimator
GE Healthcare Services Quality Form - Collimator
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Copyright 2014 General Electric Company. All rights reserved.
Collimator Replacement Verification Data
Alignments and Calibration
Test Result
Flash Download was performed and the results indicate a passing condition.
Gantry Rotation Safety Check was performed and the results indicate a passing condition.
Gantry Balance Check was performed and the results indicate a passing condition.
Z-Alignment and ISO Alignment were performed and the results indicate a passing condition.
Collimator Calibration was performed and the results indicate a passing condition.
Detailed Calibration was performed and the results indicate a passing condition.
Fast Calibration was performed and the results indicate a passing condition.
Functional System Tests was performed and the results indicate a passing condition.
Collimator Aperature test passed.
Collimator Application Position Test Passed. This test is part of the collimator functional test on the Common Service Desktop. Select all detector coverage options and filter options; click Run. When the test is done (in a minute or two), the final line in the status window should read: "Test has completed with 0 errors".
HHS data has been entered in the work sheet below (I.e. HHS work sheet).
Image Quality |  | Test Result
Image Quality test was performed and the results indicate a passing condition.
Comments:
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  | Page 2
GE Confidential

## Console
GE Healthcare Services Quality Form - Console
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Copyright 2014 General Electric Company. All rights reserved.
System Warning Labels
SCIM or GSCB warning labels:
Test Result
The Xray warning indicator is attached to the SCIM or GSCB.
The SCIM or GSCB Xray warning statement is in the appropriate language.
System Ground Check | Specification |  |  | Data | Test Result
| Min | Max | Units
Table/Gantry raceway to Console ground wire resistance. | 0 | <1.0 | ohm
System Ground Chassis Leakage test was performed and the results indicate a passing condition.
Patient Touch Leakage test was performed and the results indicate a passing condition.
Software
Console software |  |  |  | Test Result
Load From Cold performed.
All computer integration tasks were completed per instructions in the installation manual.
User Interface Language
Note: The UI is expected to be using the language appropriate for the country in which this system is installed, according to Direction 5221102-1EN. If an authorized site representative chooses another UI language, record the following information. Otherwise enter 'NA'.
Radiology Manager (or equivalent)
Date of discussion
Language selected
Appropriate or Requested language is installed.
Operator Controls / Indicators |  |  |  | Test Result
The SCIM or GSCB Xray On indicator illuminates during a scan.
The SCIM or GSCB stop scan controls terminates a scan when pressed.
Functional System Tests |  |  |  | Test Result
Fast Calibration was performed and the results indicate a passing condition.
Functional System Tests was performed and the results indicate a passing condition.
Image Quality |  |  |  | Test Result
Image Quality test was performed and the results indicate a passing condition.
Comments:
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  |  |  | Page 3
GE Confidential

## Gantry
GE Healthcare Services Quality Form - Gantry
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Copyright 2014 General Electric Company. All rights reserved.
System Warning Labels
Laser warning labels:
Note: The warning label installed must be written in the appropriate language for the Country the system is installed in.
Test Result
Laser warning labels are installed on the lower right corner of the Gantry (near the floor) and near each laser window opening in the Gantry covers.
Laser warning labels are in the appropriate language.
Gantry/Table warning labels |  |  |  | Test Result
All warning labels were installed following the System Installation Manual.
All warning labels are in the appropriate language as defined in Direction 5221102-1EN.
Gantry Anchoring
| Specifications |  |  |  | Test Result
| Min | Max | Units | Data
Gantry - Anchor type | NA | NA | NA
Gantry - Anchor supplied by | NA | NA | NA
Gantry - Length of anchor extending above the securing nut is entered in the field to the right and is within the specification limits. | 12 | 21 | mm
Gantry - Anchor torque used is entered in the field to the right and is within the specification limits. | 75 (55) | 75 (55) | N-m (lb-ft)
Gantry - Anchoring conforms to the specifications | NA | NA | NA | NA
System Interconnect Cables Continuity Test
Test Result
All cables are secured within their raceways without being damaged or pinched.
System Ground Check | Specification |  |  | Data | Test Result
| Min | Max | Units
Table/Gantry raceway to Gantry ground wire resistance. | 0 | <1.0 | ohm
Site Ground Continuity Check was performed and the results indicate a passing condition.
Patient Touch Leakage test was performed and the results indicate a passing condition.
System Chassis Leakage Test was performed and the results indicate a passing condition.
System Tests
|  |  |  | Data | Test Result
Gantry Bearing Gap was inspected and found tobe acceptable.
Interference test was performed and the results indicate a passing condition. Note: The cradle extender must be installed for this test.
Alignment Lights |  |  |  |  | Test Result
Internal axial lights are functional and aligned.
External axial lights are functional and aligned.
Coronal lights are functional and aligned.
External Alignment lights are functional (if present).
Table Cradle alignment light accuracy test passed.
Xray On Indicators |  |  |  |  | Test Result
The Gantry front cover Xray On indicator illuminates during a scan.
The Gantry rear cover Xray On indicator illuminates during a scan.
Tomographic Plane Indicator Accuracy | Specification |  |  |  | Test Result
| Min | Max | Units | Data
Internal landmark scan plan deviation indicates a passing condition. | -1.0 | +1.0 | mm
External landmark scan plan deviation indicates a passing condition. | -1.0 | +1.0 | mm
Note: If scan location S1.0 or scan location I1.0 has the darkest bar, the system meets the desired specification.
Functional System Tests |  |  |  |  | Test Result
Collimator Calibration was performed and the results indicate a passing condition.
Detailed Calibration was performed and the results indicate a passing condition.
Fast Calibration was performed and the results indicate a passing condition.
Functional System Tests was performed and the results indicate a passing condition.
Image Quality |  |  |  |  | Test Result
Image Quality test was performed and the results indicate a passing condition.
Comments:
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  |  |  |  | Page 4
GE Confidential

## HV Generator
GE Healthcare Services Quality Form - HV Generator
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Copyright 2014 General Electric Company. All rights reserved.
HV Generator Replacement Verification Data
Alignments and Calibration
Test Result
Gantry Rotation Safety Check was performed.
Gantry Balance procedure was performed.
Meter Verification performed. Enter data below.
High Voltage Feedback Resistor Verification performed. Enter data below.
Internal Scan Timer Verification performed. Enter data below.
Filament Calibration was performed and the results indicate a passing condition.
HHS data has been entered in the HHS work sheet.
Collimator Calibration was performed and the results indicate a passing condition.
Detailed Calibration was performed and the results indicate a passing condition.
Fast Calibration was performed and the results indicate a passing condition.
Image Quality |  | Test Result
Image Quality test was performed and the results indicate a passing condition.
Comments:
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  | Page 5
GE Confidential

## Tube
GE Healthcare Services Quality Form - X-ray Tube
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Copyright 2014 General Electric Company. All rights reserved.
X-ray Tube Replacement Verification Data
Alignments and Calibration
Test Result
Gantry Rotation Safety Check was performed.
Gantry Balance procedure was performed.
Tube Install Certification was performed.
Reset TnT was performed.
Z-Alignment was performed.
Cold ISO Alignment was performed.
Meter Verification performed. Enter data below.
High Voltage Feedback Resistor Verification performed. Enter data below.
Filament Calibration was performed and the results indicate a passing condition.
HHS data has been entered in the HHS work sheet.
Hot ISO alignment was performed.
Collimator Calibration was performed and the results indicate a passing condition.
Detailed Calibration was performed and the results indicate a passing condition.
Fast Calibration was performed and the results indicate a passing condition.
Image Quality |  | Test Result
Image Quality test was performed and the results indicate a passing condition.
Comments:
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  | Page 6
GE Confidential

## Table
GE Healthcare Services Quality Form - Table
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Copyright 2014 General Electric Company. All rights reserved.
Table / Gantry Alignment
| Specifications |  |  |  | Test Result
| Min | Max | Units | Data
Cradle Level - X direction | Bubble centered in finder. |  | NA
Cradle Level - Z direction | Bubble centered in finder. |  | NA
Cradle is level in all directions. | -0.1 | 0.1 | degree
Table/Gantry alignment was completed per Book 1 of the Installation Manual. | NA | NA | NA | NA
Table/Gantry alignment is within specification. | -1 | 1 | mm
Table - Anchor type | NA | NA | NA
Table - Anchor supplied by | NA | NA | NA
Table - Anchoring conforms to the specifications | NA | NA | NA | NA
Table - Length of anchor extending above the securing nut is entered in the field to the right and is within the specification limits. | 12 | 21 | mm
Table - Anchor torque used is entered in the field to the right and is within the specification limits. | 75 (55) | 75 (55) | N-m (lb-ft)
Floor type is concrete and of the thichness is entered in the field to the right. | 102 | NA | mm
Note: If the Floor type entered is 'Other', identify the floor material and thickness in this field. Otherwise enter "NA" in this field.
Floor thickness, level, and material conforms to the specifications.
System Ground Check | Specification |  |  | Data | Test Result
| Min | Max | Units
Table/Gantry raceway to Gantry ground wire resistance. | 0 | <1.0 | ohm
Site Ground Continuity Check was performed and the results indicate a passing condition.
System Chassis Leakage Test was performed and the results indicate a passing condition.
Alignment Lights
Internal axial lights are functional and aligned.
External axial lights are functional and aligned.
Coronal lights are functional and aligned.
External Alignment lights are functional (if present).
Table Cradle alignment light accuracy test passed.
Xray On Indicators
The SCIM or GSCB Xray On indicator illuminates during a scan.
The Gantry front cover Xray On indicator illuminates during a scan.
The Gantry rear cover Xray On indicator illuminates during a scan.
The Room Xray On indicator illuminates during a scan.
The Room door interlock stops a scan when the door is open (if present).
Tomographic Plane Indicator Accuracy | Specification |  |  |  | Test Result
| Min | Max | Units | Data
Internal landmark scan plan deviation indicates a passing condition. | -1.0 | +1.0 | mm
External landmark scan plan deviation indicates a passing condition. | -1.0 | +1.0 | mm
Note: If scan location S1.0 or scan location I1.0 has the darkest bar, the system meets the desired specification.
Functional System Tests |  |  |  |  | Test Result
Fast Calibration was performed and the results indicate a passing condition.
Interference test was performed and the results indicate a passing condition. Note: The cradle extender must be installed for this test.
Functional System Tests was performed and the results indicate a passing condition.
Image Quality |  |  |  |  | Test Result
Image Quality test was performed and the results indicate a passing condition.
Comments:
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  |  |  | Page 7
GE Confidential

## HHS
GE Healthcare Services Quality Form - HHS
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
HHS Data for CT Systems
Copyright 2014 General Electric Company. All rights reserved.
Note: Results of the HHS scan are saved in the following file: /usr/g/service/log/gencal.hhs_scan.report This file may be inserted into this worksheet as an alternative to entering the data in the table below.
Small Focal Spot |  | On Screen Data |  |  |  | Large Focal Spot |  | On Screen Data
kV | mA | kV | 5ms mA | Time |  | kV | mA | kV | 5ms mA | Time
80 | [5 empty] | 80
kV | mA | kV | 5ms mA | Time |  | kV | mA | kV | 5ms mA | Time
100 | [5 empty] | 100
kV | mA | kV | 5ms mA | Time |  | kV | mA | kV | 5ms mA | Time
120 | [5 empty] | 120
kV | mA | kV | 5ms mA | Time |  | kV | mA | kV | 5ms mA | Time
140 | [5 empty] | 140
Meter Verification |  |  | Specification |  | DMM Reading |  | A/D | Delta | Test Result
|  |  | Units | Max
mA Low Scale |  |  | mA | +/- 3.0
mA High Scale |  |  | mA | +/- 3.0
kV |  |  | NA | NA
HV Tank Feedback Resistor Verification | [6 empty] | Measured kV |  | On Screen kV
80
100
120
140
Scan Timer Value
DOC0755493 rev 4 |  |  |  | GE4879 - GE Healthcare Services Quality Form | [5 empty] | Page 8
GE Confidential

## RevHistory
GE Healthcare Services Quality Form - Revision History
Parent document: Optima CT660/Optima CT680/Revolution EVO Part Replacement and Service Directions
Copyright 2014 General Electric Company. All rights reserved.
Revision | Date | Author | Description of Change
1 | 2010-04-20 | Takashi Uehara | Initial release.
2 | 2014-03-11 | Takashi Uehara | Converted format from Word to Excel.
3 | 2014-06-04 | Takashi Uehara | Addition of Revolution EVO.
4 | 2014-08-21 | Takashi Uehara | Added Optima CT680 in Parent Document. Added Gantry Rotation Safety Check in Collimator, HV Geneartor and Tube replacements.
Template Information
Document | DOC0755493
Revision | 4
Title | GE4879 - GE Healthcare Services Quality Form
Footer | GE Confidential
Proprietary | Copyright 2014 General Electric Company. All rights reserved.
DOC0755493 rev 4 | GE4879 - GE Healthcare Services Quality Form |  |  | Page 9
GE Confidential