[image: image1.png]



GE Medical Systems
[image: image1.png]
Technical 

Publications
GE MEDICAL SYSTEMS
signa openspeed pRE-Order SURVEY 

REV 0
dIRECTION 2274391
DIRECTION XXXXX or “Filename” (e.g. WDtmplt.DOC)


Direction 2274391 
Revision 0

Signa OpenSpeed 

Pre-Order Survey


31- Introduction & Scope

1-1 Pre-Sale/Site Survey Scope
3
1-2 Pre-Sale/Site Survey Instructions
3
1-3 Electronically Completing This Document
3
1-4 Electronically Navigating Document Fields
3
1-5 OpenSpeed Order Processing Flowcharts
4
2- SITE LOCATION & SURVEY LOGISTICS
7
3- SURVEY QUESTIONS
7



1- Introduction & Scope
1-1 Pre-Sale/Site Survey Scope
Due to the vertical field of the Signa OpenSpeed 0.7T Magnet, the site requirements are significantly different from other GE magnets.  Therefore, accurate completion of this very basic pre-sale/site survey is critical to making sure the site qualifies for a more detailed screening survey, which must be completed prior order entry.  OpenSpeed Order Processing Flowcharts have also been included to outline the steps to be taken and the control points that have been added.

1-2 Pre-Sale/Site Survey Instructions
This Initial Pre-Order survey must be completed and given to the Local Customer Team (LCT) Installation Specialist.  The LCT Installation Specialist will then make sure a Signa OpenSpeed Site Screening Survey is completed and signed off before an order can be approved for a Signa OpenSpeed.  See Illustration 1-1, Pre-Sale/Site Survey Process Map.


Note 
This document (Direction 2274391) is the first of two steps in the site qualification process for a Signa OpenSpeed Magnet.  Completion may require input supplied by the Customer and by GE.  Do not forward this document until all questions are answered “yes”.  This helps to ensure that the Customer is aware of the basic siting requirements and has been able to identify a location that can then be surveyed in greater detail.  An order for Signa OpenSpeed will not be processed until a Signa OpenSpeed Site Screening Survey (Direction 2267841) is reviewed and approved by the LCT Installation Specialist.

1-3 Electronically Completing This Document
This document is formatted in Microsoft WORD with Text Form Fields for entry of information and Check Box Fields for selections.  

Text Fields:

· The text fill-in fields are underlined blanks.

· The text field will expand if additional space is needed.

Check Box Fields:

· The checkbox fields appear as a box.

· Clicking on blank check box will change the box to checked (X), the check box can be toggled to no X by an additional click on check box. 

1-4 Electronically Navigating Document Fields  
You can move sequentially through the document text and check box fields by depressing the following keys.

Forward: Tab, Page Up, ( key.
Backward: Shift and Tab, Page Down, ( key.
Use the side scroll bar to move to other locations in the document (i.e. Table Of Contents (page 2)) or GoTo (under Edit) to select/move to a specific page number.


1-5 OpenSpeed Order Processing Flowcharts
[image: image2.png]ON GRADE WITH MAGNETIC & RF SHIELD
TYPICAL ELEVATION

TYPICAL

8" [200mm]
SUPPORT
STRUCTURE

MAGNETIC
-+ SHIELDING

CLEAR
STRUCTURE

18" [457mm]
— MINIMUM CONCRETE

T0 BE — 1 [25mm]/

DETERMINED ISOLATION JOINT

ENGINEERING FILL ——""




[image: image3.png]ON GRADE RF SHIELD ONLY WITH 18" [457mm] SLAB
TYPICAL ELEVATION

i i

f / ]

‘ RF SCREEN / ‘

" \ |

s taS |« far \
JEaE T95S

O % & } 0O T Jr }

‘ 18" [457mm] ‘

| — MINIMUM CONCRETE |

af

a <
a a : B} .

1" [ZSmm]/ R

ISOLATION JOINT




[image: image4.png]ON GRADE WITH MAGNETIC & RF SHIELD
TYPICAL ELEVATION

TYPICAL

8" [200mm]
SUPPORT
STRUCTURE

MAGNETIC
-+ SHIELDING

CLEAR
STRUCTURE

18" [457mm]
— MINIMUM CONCRETE

T0 BE — 1 [25mm]/

DETERMINED ISOLATION JOINT

ENGINEERING FILL ——""






  

  


2- SITE LOCATION & SURVEY LOGISTICS
Customer Name:
     

Proposed Location Identifier:
     

Survey Completion Date:
     

Individual completing Survey:
     

Individual’s Title:
     

3- SURVEY QUESTIONS
These questions are the absolute minimums that must be met to pre-screen a site for Signa OpenSpeed.  All questions must be answered YES for each proposed location.  Without meeting these minimums the facility can not support the Signa OpenSpeed system.  Refer to the Signa OpenSpeed Pre-Order Survey drawings (Illustration 3-1 & 3-2) for visual representation of these siting considerations.

QUESTIONS
YES
NO

1. Is proposed site Magnet Room located on Grade (ground floor of building with Earth in contact with the floor)  with minimum 18 inch (457 mm) concrete slab below the entire Magnet Room floor? OR  If there is not 18 inch (457 mm) slab below the entire Magnet Room, can site be modified to meet this requirement?
 FORMCHECKBOX 

 FORMCHECKBOX 


2. Is there at least 9 feet (2.74 meter) vertical height in the Magnet Room without any occupancy above the Magnet Room? OR  If the area above the proposed magnet location is occupied, then is there at least 10 feet (3.05m) vertical height in the Magnet Room? (The room may require magnetic shielding.  Illustration 3-2)
 FORMCHECKBOX 

 FORMCHECKBOX 


3. Does the Magnet Room have direct access to the outside of the structure for delivery and installation of the magnet? OR  Is there a delivery path with doors/hallways opening Width x Height of at least 93.5 inches x 96.5 inches (2375 mm x 2451 mm)?
 FORMCHECKBOX 

 FORMCHECKBOX 


4. All vehicular traffic including parking areas, elevators and any other large movable metal objects are beyond at least 25ft.(7.62m) of Iso-Center of the proposed magnet location?
 FORMCHECKBOX 

 FORMCHECKBOX 


5. The Customer understands the importance of the facility meeting the siting requirements and the severe impact on system performance in failing to do so.  
(The local Installation Specialist can provide names of external vendors that can assist the customer with meeting siting requirements.)
 FORMCHECKBOX 

 FORMCHECKBOX 


Send copies of this page to Local Customer Team Installation Specialist to enable technical survey actions (Direction 2267841 Signa OpenSpeed Site Screening Survey).


Illustration 3-1 


Illustration 3-2

Revision History 
REV
DATE
AUTHOR
PRIMARY REASONS FOR CHANGE

Dev
May11, 2000
D. Balis
Development revision of document




























































Customer communicates that site is ready and Pre-Installation  Manual (Direction 2241391) Requirements are met and the site is ready for magnet delivery.

































































No





Yes





Yes





Yes





No





No





No





No





Yes





Yes





OpenSpeed Orders





If Solution or Alternate Location Can Not be Identified Order is Canceled





This Flowchart Assumes that the Pre-Order Survey(Direction 2274391) is Complete / Passed and that an Order Has Been Entered.





Local Installation Specialist Works With Customer to Define Solution or an Alternate Location and Restarts the Site Screening Process





If and Alternate Site Can Not be Identified Order is Canceled





OpenSpeed Site Screening Process Flowchart�Direction 2267841





Magnet Delivery Is Postponed To Later Date as Determined  By Re-mediation Activity.   





Customer Communicates that site is finished and meets requirements?





Customer Communicates that site is finished and meets requirements?





Installation Specialist Reviews Survey Results, Preliminary Drawings and Re-Emphasizes Pre-Installation Manual Requirements with Customer.





Installation Specialist Completes OpenSpeed Site Screening Survey


Direction 2267841





Magnet Delivered and Installation Continues





Survey Passes?





Start





Provide Customer with Signa OpenSpeed Pre-Installation  Direction 2241391 


Emphasizing Key Requirements (On-Grade, EMI and Vibration)


Provide Customer Contact Information for Questions on Siting Requirements





Installation Specialist Schedules Site Visit as Soon as Construction Allows to Continue with Site Qualification Process. 





Re-Start Site Qualification Process





List of Potential Locations is sent to LCT Installation Specialist to complete next steps.


(Direction 2267841)





LCT Works with Customer to find Alternate location(s)





One or More Locations Pass?





Proposed Site is New Construction?





Start





Order Process Is Continued





Sales Competes OpenSpeed Pre-Order Survey to Determine Possible One or More Locations (Direction 2274391)





OpenSpeed Pre-Order Survey Process Flowchart �Direction 2274391





Non OpenSpeed Orders





Order Process Difference/Control Points For Signa OpenSpeed Orders





Installation Specialist reviews Pre-Installation manual w/Customer and provides contact information for local Install Specialist and Installation Services.  Signa OpenSpeed Site Screening Survey (Direction 2267841) is completed and any outstanding concerns are discussed with Customer. 


(See Page 6)





Sales Completes Pre-Order Survey


Direction 2274391


(See Page 5)





System Transferred





System Installed





Magnet Delivered





Order Scheduled





InstallPro


Completed





Potential Signa OpenSpeed Customer Identified and Receive request for Quote.





System Transferred





System Installed





Order Delivered





Order Scheduled





InstallPro


Completed





Potential  MR Customer Identified and Receive request for Quote.





� EMBED Word.Picture.8  ���








Copyright© 2000 by General Electric Co.
Operating Documentation 


7





_1019019897.doc
[image: image1.png]ON GRADE WITH MAGNETIC & RF SHIELD
TYPICAL ELEVATION

TYPICAL

8" [200mm]
SUPPORT
STRUCTURE

MAGNETIC
-+ SHIELDING

CLEAR
STRUCTURE

18" [457mm]
— MINIMUM CONCRETE

T0 BE — 1 [25mm]/

DETERMINED ISOLATION JOINT

ENGINEERING FILL ——""








